## general informations
Version 024 Uk
| September 2015 |  |  | Modbus Memory Plane for PONSEL DIGITAL PODS
1 memory box: 2 bytes
|  | Information available |  |  | The "float" correspond to the ANSI / IEEE standard Std 754-1985 - precision (32 bits) (MSB in mind, 'big- endian')
Used for measurement parameter setting
Used for measurement:
Used for POD parameter setting:
Used for parameter calibration
System
Not in use
| Modbus address | Memory | Read/Write |  | Type | memory box number
1 | 0x0001 |  | w | Order to run a given measurement (measurements can be run simultaneously) | int | 1 | [11 empty] | M4 | M3 | M2 | M1 | MT
2 | [6 empty] | MT: |  |  |  |  | 1: runs temperature measurement
2 | [6 empty] | M1: |  |  |  |  | 1: runs parameter 1 measurement
2 | [6 empty] | M2: |  |  |  |  | 1: runs parameter 2 measurement
2 | [6 empty] | M3: |  |  |  |  | 1: runs parameter 3 measurement
2 | [6 empty] | M4: |  |  |  |  | 1: runs parameter 4 measurement
2 | 0x0002 |  | w | Restore default coefficients | int | 1 | [11 empty] | Param4 | Param3 | Param2 | Param1 | ParamTemp
3 | [6 empty] | ParamTemp |  | 1: restores default coefficients for temperature parameter (Calibration setting, Gross setting, Coef, Status, Operator name and current Calibration modification date and temporary Calibration ppv Default Coefficient)
3 | [6 empty] | Param1 |  | 1: restores default coefficients for parameter 1 (Calibration standard, Gross Value, Coefficient, Status, Operator Name and Current and Temporary Calibration Modification Date ppv Default Coefficient)
3 | [6 empty] | Param2 |  | 1: restores default coefficients for parameter 2 (Calibration standard, Gross Value, Coefficient, Status, Operator Name and Current and Temporary Calibration Modification Date ppv Default Coefficient)
3 | [6 empty] | Param3 |  | 1: restores default coefficients for parameter 3 (Calibration standard, Gross Value, Coefficient, Status, Operator Name and Current and Temporary Calibration Modification Date ppv Default Coefficient)
3 | [6 empty] | Param4 |  | 1: restores default coefficients for parameter 4 (Calibration standard, Gross Value, Coefficient, Status, Operator Name and Current and Temporary Calibration Modification Date ppv Default Coefficient)
3 | 0x0003 |  | w | Returns one of 10 latest calibrations or one of 10 sites | int | 1 |  |  |  |  | HIST | [6 empty] | PARA 4 | PARA 3 | PARA 2 | PARA 1 | PARA T
4 | [6 empty] | PARA X: | Set to '1' to recover calibration coefficient for THE parameter concerned (transfer of operator name, calibration date, site name, calibration statuses, calibration settings, gross values and etalX coefficients to current calibration and temporary calibration).
4
4 | [6 empty] | HIST: | Setting from 1 to 10. Used to identify number of calibration standard to be recovered (or site number to be recovered)
4
4
4
4 | 0x0004 |  | w | Cut of the power supply of the sensor in more of 100ms (timeto protect registers) | int | 1 | Not used at present
5 | 0x0005 |  | w | Modify address of sensor identified by serial number (using address 0) | ascii + int | 17 | 16 first int: char chain for serial number (use only with Logtek software)
22 | [6 empty] | end int: new address
22 | #NAME? |  |  | RAM Not used |  | 54
76 | 0x004B |  | w | Reset of the sensor | int | 1 | Asset by putting the value has 0x01
76 | #NAME? |  | w | Reset standard + Operator + Date of temporary calibration | int | 1 | actif with value 0x01
77 | 0x004D |  | w | Reset Bootloader | int | 1
78 | 0x004E |  | w | Download the BootLoader | int | 1 | To start with, we cross in parameter the characters of the first 3 lines of the file A43 to be loaded, to verify the compatibility of the file, then we iron all the file lines by line up to the line ox1A00
79 | 0x004F |  | w | BootLoad the program | int | 1 | The sensor passes in mode BootLoader if the wrinting value in this register is 0x5555
80
80 | 0x0050 | RAM | r | diagnostic (in %) | int | 1 | Not used at present
81 | 0x0051 | RAM |  | ………………………………………………………………………………… | int | 1
82 | 0x0052 | RAM | r | Measurement status (Generic) | int | 1 | StmTemp | Stm4: |  |  | Stm3: |  |  | Stm2: |  |  | Stm1: |  |  | StmT:
83 | [6 empty] | StmX: |  | 000: Measurement OK
83 | [8 empty] | 001: Measurement OK, but out of specifications
83 | [8 empty] | 010: Measurement OK with INFO 1
83 | [8 empty] | 011: Measurement OK with INFO 2
83 | [8 empty] | 100: Measurement impossible, out of specifications
83 | [8 empty] | 101: Measurement impossible with INFO 3
83 | [8 empty] | 110: Measurement impossible with INFO 4
83 | [8 empty] | 111: measurement under way (not yet available)
83 | [6 empty] | StmTemp |  | 1: if at least one coefficient uses a temporary calibration standard (configuration during calibration)
83 | 0x0053 | RAM | r | Temperature measurement | float | 2
85 | 0x0055 | RAM | r | Parameter 1 measurement | float | 2
87 | 0x0057 | RAM | r | Parameter 2 measurement | float | 2
89 | 0x0059 | RAM | r | Parameter 3 measurement | float | 2
91 | 0x005B | RAM | r | Parameter 4 measurement | float | 2
93 | 0x005D | RAM | r / w | Compensates Temperature returned by Master | float | 2 | Default Val: 25°C
95 | 0x005F | RAM | r / w | Compensates 1 returned by Master (external parameter for calculation of compensation) | float | 2
97 | 0x0061 | RAM | r / w | Compensates 2 returned by Master (external parameter for calculation of compensation) | float | 2
99
99 | #NAME? | RAM | r | State of peripherals of the sensor | int | 1 | PERI_15 | PERI_14 | PERI_13 | PERI_12 | PERI_11 | PERI_10 | PERI_9 | PERI_8 | PERI_7 | PERI_6 | PERI_5 | PERI_4 | PERI_3 | PERI_2 | PERI_1 | PERI_0
100 | [6 empty] | PERI_x: a 1, peripheral n°x is in accordance. A 0, there is a problem with a peripheral n°x.
100
100
100
100 | 0x0064 | RAM | r | Temperature measurement status | int | 1 | [5 empty] | StmGamme |  |  | StmME | StmHIST | StmCU | StmHF | StmTemp | Stm
101 | 0x0065 | RAM | r | Parameter 1 measurement status | int | 1 | Stm: |  | See StmX
102 | 0x0066 | RAM | r | Parameter 2 measurement status | int | 1 | StmTemp: |  | 1: if at least one parameter coefficient uses a temporary calibration standard (configuration during calibration)
103 | 0x0067 | RAM | r | Parameter 3 measurement status | int | 1 | StmHF: |  | 1: if at least one parameter coefficient is outside calibration limits
104 | 0x0068 | RAM | r | Parameter 4 measurement status | int | 1 | StmCU: |  | 1: if all parameter coefficients use default coefficient
105 | [6 empty] | StmHIST: |  | 1: if all parameter coefficients use history coefficient
105 | 0x0069 | RAM | r / w | Compensates 3 returned by Master (external parameter for calculation of compensation) | float | 2 | StmME: |  | 1: if dry weight is missing in at least one parameter coefficient (measurement under way)
107 | [6 empty] | StmGamme: |  | 001: measurement used range 1
107 | [8 empty] | 010: measurement used range 2
107 | [8 empty] | 011: measurement used range 3
107 | [8 empty] | 100: measurement used range 4
107
107 | #NAME? | RAM | r | Power supply of the sensor | float | 2 | en Volt
109 | #NAME? | RAM |  | RAM Not used |  | 2
111 | #NAME? | RAM | r | Variable of type u16 in RAM 1 | int | 1
112 | #NAME? | RAM | r | Variable of type u16 in RAM 2 | int | 1
113 | #NAME? | RAM | r | Variable of type u16 in RAM 3 | int | 1
114 | #NAME? | RAM | r | Variable of type u16 in RAM 4 | int | 1
115 | #NAME? | RAM | r | Variable of type u16 in RAM 5 | int | 1
116 | #NAME? | RAM | r | Variable of type u16 in RAM 6 | int | 1
117 | #NAME? | RAM | r | Variable of type u16 in RAM 7 | int | 1
118 | #NAME? | RAM | r | Variable of type u16 in RAM 8 | int | 1
119 | #NAME? | RAM | r | Variable of type u16 in RAM 9 | int | 1
120 | #NAME? | RAM | r | Variable of type u16 in RAM 10 | int | 1
121 | #NAME? | RAM | r | Variable of type u16 in RAM 11 | int | 1
122 | #NAME? | RAM | r | Variable of type u16 in RAM 12 | int | 1
123 | #NAME? | RAM | r | Variable of type u16 in RAM 13 | int | 1
124 | #NAME? | RAM | r | Variable of type u16 in RAM 14 | int | 1
125 | #NAME? | RAM | r | Variable of type u16 in RAM 15 | int | 1
126 | #NAME? | RAM | r | Variable of type u16 in RAM 16 | int | 1
127 | #NAME? | RAM | r | Floating variable type in RAM 1 | float | 2
129 | #NAME? | RAM | r | Floating variable type in RAM 2 | float | 2
131 | #NAME? | RAM | r | Floating variable type in RAM 3 | float | 2
133 | #NAME? | RAM | r | Floating variable type in RAM 4 | float | 2
135 | #NAME? | RAM | r | Floating variable type in RAM 5 | float | 2
137 | #NAME? | RAM | r | Floating variable type in RAM 6 | float | 2
139 | #NAME? | RAM | r | Floating variable type in RAM 7 | float | 2
141 | #NAME? | RAM | r | Floating variable type in RAM 8 | float | 2
143 | #NAME? | RAM | r | Floating variable type in RAM 9 | float | 2
145 | #NAME? | RAM | r | Floating variable type in RAM 10 | float | 2
147 | #NAME? | RAM | r | Floating variable type in RAM 11 | float | 2
149 | #NAME? | RAM | r | Floating variable type in RAM 12 | float | 2
151 | #NAME? | RAM | r | Floating variable type in RAM 13 | float | 2
153 | #NAME? | RAM | r | Floating variable type in RAM 14 | float | 2
155 | #NAME? | RAM | r | Floating variable type in RAM 15 | float | 2
157 | #NAME? | RAM | r | Floating variable type in RAM 16 | float | 2
159 | #NAME? | RAM | r | usind by the system | int | 1
160 | #NAME? | Flash / RAM | r / w | joker 1 | int | 1 | Enables to know if default parameters have been load yet (0xA5 : default parameters have already been load once at least)
161 | #NAME? | Flash / RAM | r / w | joker 2 | int | 1 | To define
162 | #NAME? | Flash / RAM | r / w | rer | int | 1 | To define
163 | 0x00A3 | Flash | r / w | Sensor address | int | 1 | from 1 to 247, decimal value
164 | 0x00A4 | Flash | r | Approx Time needed to obtain all measurements | int | 1 | in ms
165 | 0x00A5 | Flash | r / w | Temperature measurement type configuration. | int | 1 |  |  | CgDif |  |  | CgGamme |  |  | Cg3Ex | Cg2Ex | Cg1Ex | CgTEx |  |  | CgMes
166 | 0x00A6 | Flash | r / w | Param 1 measurement type configuration | int | 1 | CgMes: |  |  | 00: corrected measurement, default value
167 | 0x00A7 | Flash | r / w | Param 2 measurement type configuration | int | 1 |  |  |  | 01: uncorrected measurement 1
168 | 0x00A8 | Flash | r / w | Param 3 measurement type configuration | int | 1 |  |  |  | 10: uncorrected measurement 2
169 | 0x00A9 | Flash | r / w | Param 4 measurement type configuration | int | 1 |  |  |  | 11: gross measurement
170 | [6 empty] | CgTEX: |  |  | 1: Uses temperature returned by Master for compensation calculation. If not, Pod carries out temperature measurement and uses value for compensation.
170 | [6 empty] | Cg1EX: |  |  | 1: Pod uses Compensates 1 returned by Master for compensation calculation. If not, Pod uses default value
170 | [6 empty] | Cg2EX: |  |  | 1: Pod uses Compensates 2 returned by Master for compensation calculation. If not, Pod uses default value
170 | [6 empty] | Cg3EX: |  |  | 1: Pod uses Compensates 3 returned by Master for compensation calculation. If not, Pod uses default value
170 | [6 empty] | CgGamme: |  |  | 000: range definition is automatic, or if no range
170 | [9 empty] | 001: use of measurement range 1
170 | [9 empty] | 010: use of measurement range 2
170 | [9 empty] | 011: use of measurement range 3
170 | [9 empty] | 100: use of measurement range 4
170 | [6 empty] | CgDif: Differed calibration |  |  | 00: Default Val
170 | [9 empty] | 01: gross measurement configuration saved for future use (activation during calibration standard write process)
170 | [9 empty] | 10: warns that gross value saved will be used to calculate parameter with calibration standard (when calibration standard next required)
170
170
170 | 0x00AA | Flash | r / w | Averaging value for all parameters except temperature. | int | 1 | from 1 to 50. When value is changed, reinitialize all averaging table values. Averaging calculation excludes initialized values.
171 | 0x00AB | Flash | r / w | Measure delaying cycle counter | long | 2 | One accumulate the number of measure since Pod creation (counter reset possibility)
173 | 0x00AD | Flash | r / w | Measure cycle counter for parameter 1 | long | 2 | One accumulate the number of measure since Pod creation (counter reset possibility)
175 | 0x00AF | Flash | r / w | Measure cycle counter for parameter 2 | long | 2 | One accumulate the number of measure since Pod creation (counter reset possibility)
177 | 0x00B1 | Flash | r / w | Measure cycle counter for parameter 3 | long | 2 | One accumulate the number of measure since Pod creation (counter reset possibility)
179 | 0x00B3 | Flash | r / w | Measure cycle counter for parameter 4 | long | 2 | One accumulate the number of measure since Pod creation (counter reset possibility)
181
181 | 0x00B5 | Flash | r / w | Fixed-mount (low Power) | int | 1 | [15 empty] | CgLowP
182 | [6 empty] | CgLowP : |  |  |  |  | 0 : mode non Low power.
182 | [11 empty] | 1 : Mode Low Power
182
182 | 0x00B6 | Flash | r | Test bench status | int | 1 | [15 empty] | StBT
183 | [6 empty] | StBT = 0x0101 : Electronic card validated by the program Excel of autotest
183 | [6 empty] | StBT = 0x0202 : The electronic card was connected to the prog Excel of autotest but not VALID. Either the operator did not end the test, or the test was bad
183 | [6 empty] | StBT = 0x0303 : The electronic card was not connected to the program Excel of autotest
183
183 | #NAME? | Flash | r / w | SDI12 Sensor address | int | 1 | value in ASCII
184 | 0x00B8 | Flash | r | Primary key for the BDD | int 64bits | 4 | Low weight in head
188 | #NAME? | Flash | r / w | Configuration of the frame of the Modbus transmission (parity) | int | 1 | [14 empty] | Par
189 | [6 empty] | The answer has the command of change of parity is made with the common parity before applying the new parity.
189 | [6 empty] | with | 00 : (By default) 2 bits of stop (no parity)
189 | [7 empty] | 01 : Even parity and 1 bit of stop
189 | [7 empty] | 10: odd parity and 1 bit of stop
189
189 | #NAME? | Flash | r | Memory location not used |  | 2
191 | 0x00BF | Flash | r | Used by the system | int | 1
192
192 | 0x00C0 | Flash | r | Calibration status of the slope of the temperature (current calibration) | int | 1 | SteE : State the point of calibration |  |  |  |  | 00 : Calibration point OK (calibration is not out of range)
193 | 0x00C1 | Flash | r | Status of calibration coefficient 1 (current calibration) | int | 1 | [5 empty] | 01 : Out of range calibration point
194 | 0x00C2 | Flash | r | Status of calibration coefficient 2 (current calibration) | int | 1
195 | 0x00C3 | Flash | r | Status of calibration coefficient 3 (current calibration) | int | 1 | SteCB |  |  |  |  | 1: Out of range low coefficient
196 | 0x00C4 | Flash | r | Status of calibration coefficient 4 (current calibration) | int | 1 | SteCH |  |  |  |  | 1: Out of range high coefficient
197 | 0x00C5 | Flash | r | Status of calibration coefficient 5 (current calibration) | int | 1 | SteCU |  |  |  |  | 1 : we use the factory settings (coefficient)
198 | 0x00C6 | Flash | r | Status of calibration coefficient 6 (current calibration) | int | 1 | SteHIST |  |  |  |  | 1 : we use the coefficient of a history
199 | 0x00C7 | Flash | r | Status of calibration coefficient 7 (current calibration) | int | 1 | SteME |  |  |  |  | 1 : measurement recorded (deferred pending)
200 | 0x00C8 | Flash | r | Status of calibration coefficient 8 (current calibration) | int | 1 | SteRE |  |  |  |  | 1 : The calibration of this coefficient was performed
201 | 0x00C9 | Flash | r | Status of calibration coefficient 9 (current calibration) | int | 1
202 | 0x00CA | Flash | r | Status of calibration coefficient 10 (current calibration) | int | 1
203 | 0x00CB | Flash | r | Status of calibration coefficient 11 (current calibration) | int | 1
204 | 0x00CC | Flash | r | Status of calibration coefficient 12 (current calibration) | int | 1
205 | 0x00CD | Flash | r | Status of calibration coefficient 13 (current calibration) | int | 1
206 | 0x00CE | Flash | r | Status of calibration coefficient 14 (current calibration) | int | 1
207 | 0x00CF | Flash | r | Status of calibration coefficient 15 (current calibration) | int | 1
208 | 0x00D0 | Flash | r | Status of calibration coefficient 16 (current calibration) | int | 1
209 | 0x00D1 |  |  |  | int
209
209 | 0x00D2 | Flash | r | The temperature Offset (current calibration) | float | 2 | Coefficient used to calculate when applying measurement
211 | 0x00D4 | Flash | r | Slope of temperature (current calibration) | float | 2 | Coefficient used to calculate when applying measurement
213 | 0x00D6 | Flash | r | Correction Coefficient 1 (current calibration) | float | 2 | Coefficient used to calculate when applying measurement
215 | 0x00D8 | Flash | r | Correction Coefficient 2 (current calibration) | float | 2 | Coefficient used to calculate when applying measurement
217 | 0x00DA | Flash | r | Correction Coefficient 3 (current calibration) | float | 2 | Coefficient used to calculate when applying measurement
219 | 0x00DC | Flash | r | Correction Coefficient 4 (current calibration) | float | 2 | Coefficient used to calculate when applying measurement
221 | 0x00DE | Flash | r | Correction Coefficient 5 (current calibration) | float | 2 | Coefficient used to calculate when applying measurement
223 | 0x00E0 | Flash | r | Correction Coefficient 6 (current calibration) | float | 2 | Coefficient used to calculate when applying measurement
225 | 0x00E2 | Flash | r | Correction Coefficient 7 (current calibration) | float | 2 | Coefficient used to calculate when applying measurement
227 | 0x00E4 | Flash | r | Correction Coefficient 8 (current calibration) | float | 2 | Coefficient used to calculate when applying measurement
229 | 0x00E6 | Flash | r | Correction Coefficient 9 (current calibration) | float | 2 | Coefficient used to calculate when applying measurement
231 | 0x00E8 | Flash | r | Correction Coefficient 10 (current calibration) | float | 2 | Coefficient used to calculate when applying measurement
233 | 0x00EA | Flash | r | Correction Coefficient 11 (current calibration) | float | 2 | Coefficient used to calculate when applying measurement
235 | 0x00EC | Flash | r | Correction Coefficient 12 (current calibration) | float | 2 | Coefficient used to calculate when applying measurement
237 | 0x00EE | Flash | r | Correction Coefficient 13 (current calibration) | float | 2 | Coefficient used to calculate when applying measurement
239 | 0x00F0 | Flash | r | Correction Coefficient 14 (current calibration) | float | 2 | Coefficient used to calculate when applying measurement
241 | 0x00F2 | Flash | r | Correction Coefficient 15 (current calibration) | float | 2 | Coefficient used to calculate when applying measurement
243 | 0x00F4 | Flash | r | Correction Coefficient 16 (current calibration) | float | 2 | Coefficient used to calculate when applying measurement
245
245 | 0x00F6 | Flash | r | Name of the operator who was calibrated temperature (current calibration) | ascii | 8
253 | 0x00FE | Flash | r | Temperature calibration date (current calibration) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
261 | 0x0106 | Flash | r | Name of the operator who calibrates the parameter 1 (current calibration) | ascii | 8
269 | 0x010E | Flash | r | Calibration date of the parameter 1 (current calibration) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
277 | 0x0116 | Flash | r | Name of the operator who calibrates the parameter 2 (current calibration) | ascii | 8
285 | 0x011E | Flash | r | Calibration date of the parameter 2 (current calibration) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
293 | 0x0126 | Flash | r | Name of the operator who calibrates the parameter 3 (current calibration) | ascii | 8
301 | 0x012E | Flash | r | Calibration date of the parameter 3 (current calibration) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
309 | 0x0136 | Flash | r | Name of the operator who calibrates the parameter 4 (current calibration) | ascii | 8
317 | 0x013E | Flash | r | Calibration date of the parameter 4 (current calibration) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
325 | 0x0146 | Flash | r | Site name (current calibration) | ascii | 4 | used onlmy for the Nephelo in mg/L and the MES
329
329
329
329 | 0x014A |  |  |  | long | 2
331
331
331 | 0x014C | Flash | r / w | List of temporary coefficients to be used in measurement calculation | long | 2 | Coef14 | Coef13 | Coef12 | Coef11 | Coef10 | Coef9 | Coef8 | Coef7 | Coef6 | Coef5 | Coef4 | Coef3 | Coef2 | Coef1 | PenteTemp | OffTemp | LSB word
333 | [20 empty] | Coef16 | Coef15 | MSB word
333 | [6 empty] | CoefX: |  | 1: Pod will use temporary coefficient value for measurement calculation instead of current coefficient; this function only used during calibration. Example: offset already calculated used to display measurement during second calibration (slope)
333
333 | 0x014E | Flash | r / w | Site number | int | 1 | Identifies site to be used for temporary calibration transfer (value from 1 to 10)
334
512 | 0x0200 | Flash | r / w | write calibration standard 1 for temperature | float | 2 | During write process, runs coefficient calculation and stores value in temporary calibration
514 | 0x0202 | Flash | r / w | write calibration standard 2 for temperature | float | 2 | During write process, runs coefficient calculation and stores value in temporary calibration
516 | 0x0204 | Flash | r / w | write calibration standard 1 | float | 2 | During write process, runs coefficient calculation and stores value in temporary calibration
518 | 0x0206 | Flash | r / w | write calibration standard 2 | float | 2 | During write process, runs coefficient calculation and stores value in temporary calibration
520 | 0x0208 | Flash | r / w | write calibration standard 3 | float | 2 | During write process, runs coefficient calculation and stores value in temporary calibration
522 | 0x020A | Flash | r / w | write calibration standard 4 | float | 2 | During write process, runs coefficient calculation and stores value in temporary calibration
524 | 0x020C | Flash | r / w | write calibration standard 5 | float | 2 | During write process, runs coefficient calculation and stores value in temporary calibration
526 | 0x020E | Flash | r / w | write calibration standard 6 | float | 2 | During write process, runs coefficient calculation and stores value in temporary calibration
528 | 0x0210 | Flash | r / w | write calibration standard 7 | float | 2 | During write process, runs coefficient calculation and stores value in temporary calibration
530 | 0x0212 | Flash | r / w | write calibration standard 8 | float | 2 | During write process, runs coefficient calculation and stores value in temporary calibration
532 | 0x0214 | Flash | r / w | write calibration standard 9 | float | 2 | During write process, runs coefficient calculation and stores value in temporary calibration
534 | 0x0216 | Flash | r / w | write calibration standard 10 | float | 2 | During write process, runs coefficient calculation and stores value in temporary calibration
536 | 0x0218 | Flash | r / w | write calibration standard 11 | float | 2 | During write process, runs coefficient calculation and stores value in temporary calibration
538 | 0x021A | Flash | r / w | write calibration standard 12 | float | 2 | During write process, runs coefficient calculation and stores value in temporary calibration
540 | 0x021C | Flash | r / w | write calibration standard 13 | float | 2 | During write process, runs coefficient calculation and stores value in temporary calibration
542 | 0x021E | Flash | r / w | write calibration standard 14 | float | 2 | During write process, runs coefficient calculation and stores value in temporary calibration
544 | 0x0220 | Flash | r / w | write calibration standard 15 | float | 2 | During write process, runs coefficient calculation and stores value in temporary calibration
546 | 0x0222 | Flash | r / w | write calibration standard 16 | float | 2 | During write process, runs coefficient calculation and stores value in temporary calibration
548
548 | 0x0224 | Flash | r | Offset Gross Value of the temp (temporary calibration) | float | 2
550 | 0x0226 | Flash | r | Gain Gross value of the temp (temporary calibration) | float | 2
552 | 0x0228 | Flash | r | Gross value 1 (temporary calibration) | float | 2
554 | 0x022A | Flash | r | Gross value 2 (temporary calibration) | float | 2
556 | 0x022C | Flash | r | Gross value 3 (temporary calibration) | float | 2
558 | 0x022E | Flash | r | Gross value 4 (temporary calibration) | float | 2
560 | 0x0230 | Flash | r | Gross value 5 (temporary calibration) | float | 2
562 | 0x0232 | Flash | r | Gross value 6 (temporary calibration) | float | 2
564 | 0x0234 | Flash | r | Gross value 7 (temporary calibration) | float | 2
566 | 0x0236 | Flash | r | Gross value 8 (temporary calibration) | float | 2
568 | 0x0238 | Flash | r | Gross value 9 (temporary calibration) | float | 2
570 | 0x023A | Flash | r | Gross value 10 (temporary calibration) | float | 2
572 | 0x023C | Flash | r | Gross value 11 (temporary calibration) | float | 2
574 | 0x023E | Flash | r | Gross value 12 (temporary calibration) | float | 2
576 | 0x0240 | Flash | r | Gross value 13 (temporary calibration) | float | 2
578 | 0x0242 | Flash | r | Gross value 14 (temporary calibration) | float | 2
580 | 0x0244 | Flash | r | Gross value 15 (temporary calibration) | float | 2
582 | 0x0246 | Flash | r | Gross value 16 (temporary calibration) | float | 2
584
584 | 0x0248 | Flash | r | Offset of the temp (temporary calibration) | float | 2
586 | 0x024A | Flash | r | Gain of the temp (temporary calibration) | float | 2
588 | 0x024C | Flash | r | Coefficient of correction 1 (temporary calibration) | float | 2
590 | 0x024E | Flash | r | Coefficient of correction 2 (temporary calibration) | float | 2
592 | 0x0250 | Flash | r | Coefficient of correction 3 (temporary calibration) | float | 2
594 | 0x0252 | Flash | r | Coefficient of correction 4 (temporary calibration) | float | 2
596 | 0x0254 | Flash | r | Coefficient of correction 5 (temporary calibration) | float | 2
598 | 0x0256 | Flash | r | Coefficient of correction 6 (temporary calibration) | float | 2
600 | 0x0258 | Flash | r | Coefficient of correction 7 (temporary calibration) | float | 2
602 | 0x025A | Flash | r | Coefficient of correction 8 (temporary calibration) | float | 2
604 | 0x025C | Flash | r | Coefficient of correction 9 (temporary calibration) | float | 2
606 | 0x025E | Flash | r | Coefficient of correction 10 (temporary calibration) | float | 2
608 | 0x0260 | Flash | r | Coefficient of correction 11 (temporary calibration) | float | 2
610 | 0x0262 | Flash | r | Coefficient of correction 12 (temporary calibration) | float | 2
612 | 0x0264 | Flash | r | Coefficient of correction 13 (temporary calibration) | float | 2
614 | 0x0266 | Flash | r | Coefficient of correction 14 (temporary calibration) | float | 2
616 | 0x0268 | Flash | r | Coefficient of correction 15 (temporary calibration) | float | 2
618 | 0x026A | Flash | r | Coefficient of correction 16 (temporary calibration) | float | 2
620
620 | 0x026C | Flash | r | Status of the calibration of the temperature offset (temporary calibration) | int | 1 | [7 empty] | SteRE |  | SteME | SteHIST | SteCU | SteCH | SteCB | SteE
621 | 0x026D | Flash | r | Status of the calibration of the temperature Gain (temporary calibration) | int | 1 | SteE : State of the calibration point |  |  |  |  | 00 : calibration point OK (the calibration is not out of range)
622 | 0x026E | Flash | r | Status of calibration of coefficient 1 (temporary calibration) | int | 1 | [5 empty] | 01 :calibration point out of range
623 | 0x026F | Flash | r | Status of calibration of coefficient 2 (temporary calibration) | int | 1
624 | 0x0270 | Flash | r | Status of calibration of coefficient 3 (temporary calibration) | int | 1 | SteCB |  |  |  |  | 1: coef out of range low
625 | 0x0271 | Flash | r | Status of calibration of coefficient 4 (temporary calibration) | int | 1 | SteCH |  |  |  |  | 1: coef out of range height
626 | 0x0272 | Flash | r | Status of calibration of coefficient 5 (temporary calibration) | int | 1 | SteCU |  |  |  |  | 1 : using of factory coef
627 | 0x0273 | Flash | r | Status of calibration of coefficient 6 (temporary calibration) | int | 1 | SteHIST |  |  |  |  | 1 : using of historical coef
628 | 0x0274 | Flash | r | Status of calibration of coefficient 7 (temporary calibration) | int | 1 | SteME |  |  |  |  | 1 : recorded measure (differed process)
629 | 0x0275 | Flash | r | Status of calibration of coefficient 8 (temporary calibration) | int | 1 | SteRE |  |  |  |  | 1 : in this historical, the calibration of this coef was realized
630 | 0x0276 | Flash | r | Status of calibration of coefficient 9 (temporary calibration) | int | 1
631 | 0x0277 | Flash | r | Status of calibration of coefficient 10 (temporary calibration) | int | 1
632 | 0x0278 | Flash | r | Status of calibration of coefficient 11 (temporary calibration) | int | 1
633 | 0x0279 | Flash | r | Status of calibration of coefficient 12 (temporary calibration) | int | 1
634 | 0x027A | Flash | r | Status of calibration of coefficient 13 (temporary calibration) | int | 1
635 | 0x027B | Flash | r | Status of calibration of coefficient 14 (temporary calibration) | int | 1
636 | 0x027C | Flash | r | Status of calibration of coefficient 15 (temporary calibration) | int | 1
637 | 0x027D | Flash | r | Status of calibration of coefficient 16 (temporary calibration) | int | 1
638
638 | 0x027E | Flash | w | Name of operator who calibrated temporary calibration standard | ascii | 8
646 | 0x0286 | Flash | w | Date of temperature calibration (temporary calibration) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
654 | [6 empty] | Validates calibration. (empties "List of coefficients to be calibrated" and "List of temporary coefficients to be used in measurement calculation" registries) If parameter is managed historically: calibration standards, gross values, coefficients, calibration statuses, operator name and calibration date for the calibrated parameter, are transferred from temporary calibration standard to current calibration standard and etal1, and history of former calibrations is set back one (etal10 disappears). If parameter is managed on site: calibration standards, gross values, coefficients, calibration statuses, operator name and calibration date for the calibrated parameter, are transferred from temporary calibration standard to current calibration standard and to the calibration standard for the site number.
654
654
654
654
654 | 0x028E | Flash | w | Name of operator who calibrated parameter 1 (temporary calibration standard) | ascii | 8
662 | 0x0296 | Flash | w | Date of calibration for parameter 1 (temporary calibration standard) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
670 | [6 empty] | Validates calibration. (empties "List of coefficients to be calibrated" and "List of temporary coefficients to be used in measurement calculation" registries) If parameter is managed historically: calibration standards, gross values, coefficients, calibration statuses, operator name and calibration date for the calibrated parameter, are transferred from temporary calibration standard to current calibration standard and etal1, and history of former calibrations is set back one (etal10 disappears). If parameter is managed on site: calibration standards, gross values, coefficients, calibration statuses, operator name and calibration date for the calibrated parameter, are transferred from temporary calibration standard to current calibration standard and to the calibration standard for the site number.
670
670
670
670
670 | 0x029E | Flash | w | Name of operator who calibrated parameter 2 (temporary calibration standard) | ascii | 8
678 | 0x02A6 | Flash | w | Date of calibration for parameter 2 (temporary calibration standard) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
686 | [6 empty] | Validates calibration. (empties "List of coefficients to be calibrated" and "List of temporary coefficients to be used in measurement calculation" registries) If parameter is managed historically: calibration standards, gross values, coefficients, calibration statuses, operator name and calibration date for the calibrated parameter, are transferred from temporary calibration standard to current calibration standard and etal1, and history of former calibrations is set back one (etal10 disappears). If parameter is managed on site: calibration standards, gross values, coefficients, calibration statuses, operator name and calibration date for the calibrated parameter, are transferred from temporary calibration standard to current calibration standard and to the calibration standard for the site number.
686
686
686
686
686 | 0x02AE | Flash | w | Name of operator who calibrated parameter 3 (temporary calibration standard) | ascii | 8
694 | 0x02B6 | Flash | w | Date of calibration for parameter 3 (temporary calibration standard) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
702 | [6 empty] | Validates calibration. (empties "List of coefficients to be calibrated" and "List of temporary coefficients to be used in measurement calculation" registries) If parameter is managed historically: calibration standards, gross values, coefficients, calibration statuses, operator name and calibration date for the calibrated parameter, are transferred from temporary calibration standard to current calibration standard and etal1, and history of former calibrations is set back one (etal10 disappears). If parameter is managed on site: calibration standards, gross values, coefficients, calibration statuses, operator name and calibration date for the calibrated parameter, are transferred from temporary calibration standard to current calibration standard and to the calibration standard for the site number.
702
702
702
702
702 | 0x02BE | Flash | w | Name of operator who calibrated parameter 4 (temporary calibration standard) | ascii | 8
710 | 0x02C6 | Flash | w | Date of calibration for parameter 4 (temporary calibration standard) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
718 | [6 empty] | Validates calibration. (empties "List of coefficients to be calibrated" and "List of temporary coefficients to be used in measurement calculation" registries) If parameter is managed historically: calibration standards, gross values, coefficients, calibration statuses, operator name and calibration date for the calibrated parameter, are transferred from temporary calibration standard to current calibration standard and etal1, and history of former calibrations is set back one (etal10 disappears). If parameter is managed on site: calibration standards, gross values, coefficients, calibration statuses, operator name and calibration date for the calibrated parameter, are transferred from temporary calibration standard to current calibration standard and to the calibration standard for the site number.
718
718
718
718
718 | 0x02CE | Flash | w | Site Name (temporary calibration standard) | ascii | 4 | used only for Nephelo in mg/L and for MES
722
722 | #NAME? | Flash | r/w | Memory space reserved for integrators. | - | 5 | The 10 bytes are at 0x00 ex-factory and for recalibration factory. Integrators can use them in int, float, ascii, ….
727
727 | 0x02D7 |  |  | Memory location not used |  | 41
768 | 0x0300 | Flash | r | Value of the temperature standard 1 (cal1) | float | 2
770 | 0x0302 | Flash | r | Value of the temperature standard 2 (cal1) | float | 2
772 | 0x0304 | Flash | r | Value of standard 1 (cal 1) | float | 2
774 | 0x0306 | Flash | r | Value of standard 2 (cal 1) | float | 2
776 | 0x0308 | Flash | r | Value of standard 3 (cal 1) | float | 2
778 | 0x030A | Flash | r | Value of standard 4 (cal 1) | float | 2
780 | 0x030C | Flash | r | Value of standard 5 (cal 1) | float | 2
782 | 0x030E | Flash | r | Value of standard 6 (cal 1) | float | 2
784 | 0x0310 | Flash | r | Value of standard 7 (cal 1) | float | 2
786 | 0x0312 | Flash | r | Value of standard 8 (cal 1) | float | 2
788 | 0x0314 | Flash | r | Value of standard 9 (cal 1) | float | 2
790 | 0x0316 | Flash | r | Value of standard 10 (cal 1) | float | 2
792 | 0x0318 | Flash | r | Value of standard 11 (cal 1) | float | 2
794 | 0x031A | Flash | r | Value of standard 12 (cal 1) | float | 2
796 | 0x031C | Flash | r | Value of standard 13 (cal 1) | float | 2
798 | 0x031E | Flash | r | Value of standard 14 (cal 1) | float | 2
800 | 0x0320 | Flash | r | Value of standard 15 (cal 1) | float | 2
802 | 0x0322 | Flash | r | Value of standard 16 (cal 1) | float | 2
804
804 | 0x0324 | Flash | r | Gross value Offset of the temp (cal 1) | float | 2
806 | 0x0326 | Flash | r | Gross value gain of the temp (cal 1) | float | 2
808 | 0x0328 | Flash | r | Gross value 1 (cal 1) | float | 2
810 | 0x032A | Flash | r | Gross value 2 (cal 1) | float | 2
812 | 0x032C | Flash | r | Gross value 3 (cal 1) | float | 2
814 | 0x032E | Flash | r | Gross value 4 (cal 1) | float | 2
816 | 0x0330 | Flash | r | Gross value 5 (cal 1) | float | 2
818 | 0x0332 | Flash | r | Gross value 6 (cal 1) | float | 2
820 | 0x0334 | Flash | r | Gross value 7 (cal 1) | float | 2
822 | 0x0336 | Flash | r | Gross value 8 (cal 1) | float | 2
824 | 0x0338 | Flash | r | Gross value 9 (cal 1) | float | 2
826 | 0x033A | Flash | r | Gross value 10 (cal 1) | float | 2
828 | 0x033C | Flash | r | Gross value 11 (cal 1) | float | 2
830 | 0x033E | Flash | r | Gross value 12 (cal 1) | float | 2
832 | 0x0340 | Flash | r | Gross value 13 (cal 1) | float | 2
834 | 0x0342 | Flash | r | Gross value 14 (cal 1) | float | 2
836 | 0x0344 | Flash | r | Gross value 15 (cal 1) | float | 2
838 | 0x0346 | Flash | r | Gross value 16 (cal 1) | float | 2
840
840 | 0x0348 | Flash | r | Offset of the temp (cal 1) | float | 2
842 | 0x034A | Flash | r | Gain of the temp (cal 1) | float | 2
844 | 0x034C | Flash | r | Coefficient of correction 1 (cal 1) | float | 2
846 | 0x034E | Flash | r | Coefficient of correction 2 (cal 1) | float | 2
848 | 0x0350 | Flash | r | Coefficient of correction 3 (cal 1) | float | 2
850 | 0x0352 | Flash | r | Coefficient of correction 4 (cal 1) | float | 2
852 | 0x0354 | Flash | r | Coefficient of correction 5 (cal 1) | float | 2
854 | 0x0356 | Flash | r | Coefficient of correction 6 (cal 1) | float | 2
856 | 0x0358 | Flash | r | Coefficient of correction 7 (cal 1) | float | 2
858 | 0x035A | Flash | r | Coefficient of correction 8 (cal 1) | float | 2
860 | 0x035C | Flash | r | Coefficient of correction 9 (cal 1) | float | 2
862 | 0x035E | Flash | r | Coefficient of correction 10 (cal 1) | float | 2
864 | 0x0360 | Flash | r | Coefficient of correction 11 (cal 1) | float | 2
866 | 0x0362 | Flash | r | Coefficient of correction 12 (cal 1) | float | 2
868 | 0x0364 | Flash | r | Coefficient of correction 13 (cal 1) | float | 2
870 | 0x0366 | Flash | r | Coefficient of correction 14 (cal 1) | float | 2
872 | 0x0368 | Flash | r | Coefficient of correction 15 (cal 1) | float | 2
874 | 0x036A | Flash | r | Coefficient of correction 16 (cal 1) | float | 2
876
876 | 0x036C | Flash | r | Status of the calibration of the temperature offset (cal 1) | int | 1 | [7 empty] | SteRE |  | SteME | SteHIST | SteCU | SteCH | SteCB | SteE
877 | 0x036D | Flash | r | Status of the calibration of the temperature gain (cal 1) | int | 1 | SteE : State of the calibration point |  |  |  |  | 00 : calibration point OK (the calibration is not out of range)
878 | 0x036E | Flash | r | Status of calibration of coefficient 1 (cal 1) | int | 1 | [5 empty] | 01 : calibration point out of range
879 | 0x036F | Flash | r | Status of calibration of coefficient 2 (cal 1) | int | 1
880 | 0x0370 | Flash | r | Status of calibration of coefficient 3 (cal 1) | int | 1 | SteCB |  |  |  |  | 1: coef out of range low
881 | 0x0371 | Flash | r | Status of calibration of coefficient 4 (cal 1) | int | 1 | SteCH |  |  |  |  | 1: coef out of range height
882 | 0x0372 | Flash | r | Status of calibration of coefficient 5 (cal 1) | int | 1 | SteCU |  |  |  |  | 1 : using of factory coef
883 | 0x0373 | Flash | r | Status of calibration of coefficient 6 (cal 1) | int | 1 | SteHIST |  |  |  |  | 1 : using of historical coef
884 | 0x0374 | Flash | r | Status of calibration of coefficient 7 (cal 1) | int | 1 | SteME |  |  |  |  | 1 : recorded measure (differed process)
885 | 0x0375 | Flash | r | Status of calibration of coefficient 8 (cal 1) | int | 1 | SteRE |  |  |  |  | 1 : in this historical, the calibration of this coef was realized
886 | 0x0376 | Flash | r | Status of calibration of coefficient 9 (cal 1) | int | 1
887 | 0x0377 | Flash | r | Status of calibration of coefficient 10 (cal 1) | int | 1
888 | 0x0378 | Flash | r | Status of calibration of coefficient 11 (cal 1) | int | 1
889 | 0x0379 | Flash | r | Status of calibration of coefficient 12 (cal 1) | int | 1
890 | 0x037A | Flash | r | Status of calibration of coefficient 13 (cal 1) | int | 1
891 | 0x037B | Flash | r | Status of calibration of coefficient 14 (cal 1) | int | 1
892 | 0x037C | Flash | r | Status of calibration of coefficient 15 (cal 1) | int | 1
893 | 0x037D | Flash | r | Status of calibration of coefficient 16 (cal 1) | int | 1
894
894 | 0x037E | Flash | r | Name of the operator which calibrated the temperature (cal 1) | ascii | 8
902 | 0x0386 | Flash | r | Date of the calibration of the temperature (cal 1) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
910 | 0x038E | Flash | r | Name of the operator which calibrated the parameter 1 (cal 1) | ascii | 8
918 | 0x0396 | Flash | r | Date of the calibration of the parameter 1 (cal 1) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
926 | 0x039E | Flash | r | Name of the operator which calibrated the parameter 2 (cal 1) | ascii | 8
934 | 0x03A6 | Flash | r | Date of the calibration of the parameter 2 (cal 1) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
942 | 0x03AE | Flash | r | Name of the operator which calibrated the parameter 3 (cal 1) | ascii | 8
950 | 0x03B6 | Flash | r | Date of the calibration of the parameter 3 (cal 1) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
958 | 0x03BE | Flash | r | Name of the operator which calibrated the parameter 4 (cal 1) | ascii | 8
966 | 0x03C6 | Flash | r | Date of the calibration of the parameter 4 (cal 1) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
974 | 0x03CE | Flash | r | Name of the site (cal 1) | ascii | 4 | only using for the turbidity sensor in mg/L and the SS sensor
978 |  |  |  | Memory location not used |  | 46
1024 | 0x0400 | Flash | r | Value of the temperature standard 1 (cal2) | float | 2
1026 | 0x0402 | Flash | r | Value of the temperature standard 2 (cal2) | float | 2
1028 | 0x0404 | Flash | r | Value of standard 1 (cal 2) | float | 2
1030 | 0x0406 | Flash | r | Value of standard 2 (cal 2) | float | 2
1032 | 0x0408 | Flash | r | Value of standard 3 (cal 2) | float | 2
1034 | 0x040A | Flash | r | Value of standard 4 (cal 2) | float | 2
1036 | 0x040C | Flash | r | Value of standard 5 (cal 2) | float | 2
1038 | 0x040E | Flash | r | Value of standard 6 (cal 2) | float | 2
1040 | 0x0410 | Flash | r | Value of standard 7 (cal 2) | float | 2
1042 | 0x0412 | Flash | r | Value of standard 8 (cal 2) | float | 2
1044 | 0x0414 | Flash | r | Value of standard 9 (cal 2) | float | 2
1046 | 0x0416 | Flash | r | Value of standard 10 (cal 2) | float | 2
1048 | 0x0418 | Flash | r | Value of standard 11 (cal 2) | float | 2
1050 | 0x041A | Flash | r | Value of standard 12 (cal 2) | float | 2
1052 | 0x041C | Flash | r | Value of standard 13 (cal 2) | float | 2
1054 | 0x041E | Flash | r | Value of standard 14 (cal 2) | float | 2
1056 | 0x0420 | Flash | r | Value of standard 15 (cal 2) | float | 2
1058 | 0x0422 | Flash | r | Value of standard 16 (cal 2) | float | 2
1060
1060 | 0x0424 | Flash | r | Gross value Offset of the temp (cal 2) | float | 2
1062 | 0x0426 | Flash | r | Gross value gain of the temp (cal 2) | float | 2
1064 | 0x0428 | Flash | r | Gross value 1 (cal 2) | float | 2
1066 | 0x042A | Flash | r | Gross value 2 (cal 2) | float | 2
1068 | 0x042C | Flash | r | Gross value 3 (cal 2) | float | 2
1070 | 0x042E | Flash | r | Gross value 4 (cal 2) | float | 2
1072 | 0x0430 | Flash | r | Gross value 5 (cal 2) | float | 2
1074 | 0x0432 | Flash | r | Gross value 6 (cal 2) | float | 2
1076 | 0x0434 | Flash | r | Gross value 7 (cal 2) | float | 2
1078 | 0x0436 | Flash | r | Gross value 8 (cal 2) | float | 2
1080 | 0x0438 | Flash | r | Gross value 9 (cal 2) | float | 2
1082 | 0x043A | Flash | r | Gross value 10 (cal 2) | float | 2
1084 | 0x043C | Flash | r | Gross value 11 (cal 2) | float | 2
1086 | 0x043E | Flash | r | Gross value 12 (cal 2) | float | 2
1088 | 0x0440 | Flash | r | Gross value 13 (cal 2) | float | 2
1090 | 0x0442 | Flash | r | Gross value 14 (cal 2) | float | 2
1092 | 0x0444 | Flash | r | Gross value 15 (cal 2) | float | 2
1094 | 0x0446 | Flash | r | Gross value 16 (cal 2) | float | 2
1096
1096 | 0x0448 | Flash | r | Offset of the temp (cal 2) | float | 2
1098 | 0x044A | Flash | r | Gain of the temp (cal 2) | float | 2
1100 | 0x044C | Flash | r | Coefficient of correction 1 (cal 2) | float | 2
1102 | 0x044E | Flash | r | Coefficient of correction 2 (cal 2) | float | 2
1104 | 0x0450 | Flash | r | Coefficient of correction 3 (cal 2) | float | 2
1106 | 0x0452 | Flash | r | Coefficient of correction 4 (cal 2) | float | 2
1108 | 0x0454 | Flash | r | Coefficient of correction 5 (cal 2) | float | 2
1110 | 0x0456 | Flash | r | Coefficient of correction 6 (cal 2) | float | 2
1112 | 0x0458 | Flash | r | Coefficient of correction 7 (cal 2) | float | 2
1114 | 0x045A | Flash | r | Coefficient of correction 8 (cal 2) | float | 2
1116 | 0x045C | Flash | r | Coefficient of correction 9 (cal 2) | float | 2
1118 | 0x045E | Flash | r | Coefficient of correction 10 (cal 2) | float | 2
1120 | 0x0460 | Flash | r | Coefficient of correction 11 (cal 2) | float | 2
1122 | 0x0462 | Flash | r | Coefficient of correction 12 (cal 2) | float | 2
1124 | 0x0464 | Flash | r | Coefficient of correction 13 (cal 2) | float | 2
1126 | 0x0466 | Flash | r | Coefficient of correction 14 (cal 2) | float | 2
1128 | 0x0468 | Flash | r | Coefficient of correction 15 (cal 2) | float | 2
1130 | 0x046A | Flash | r | Coefficient of correction 16 (cal 2) | float | 2
1132
1132 | 0x046C | Flash | r | Status of the calibration of the temperature offset (cal 2) | int | 1 | [7 empty] | SteRE |  | SteME | SteHIST | SteCU | SteCH | SteCB | SteE
1133 | 0x046D | Flash | r | Status of the calibration of the temperature gain (cal 2) | int | 1 | SteE : State of the calibration point |  |  |  |  | 00 : calibration point OK (the calibration is not out of range)
1134 | 0x046E | Flash | r | Status of calibration of coefficient 1 (cal 2) | int | 1 | [5 empty] | 01 : calibration point out of range
1135 | 0x046F | Flash | r | Status of calibration of coefficient 2 (cal 2) | int | 1
1136 | 0x0470 | Flash | r | Status of calibration of coefficient 3 (cal 2) | int | 1 | SteCB |  |  |  |  | 1: coef out of range low
1137 | 0x0471 | Flash | r | Status of calibration of coefficient 4 (cal 2) | int | 1 | SteCH |  |  |  |  | 1: coef out of range height
1138 | 0x0472 | Flash | r | Status of calibration of coefficient 5 (cal 2) | int | 1 | SteCU |  |  |  |  | 1 : using of factory coef
1139 | 0x0473 | Flash | r | Status of calibration of coefficient 6 (cal 2) | int | 1 | SteHIST |  |  |  |  | 1 : using of historical coef
1140 | 0x0474 | Flash | r | Status of calibration of coefficient 7 (cal 2) | int | 1 | SteME |  |  |  |  | 1 : recorded measure (differed process)
1141 | 0x0475 | Flash | r | Status of calibration of coefficient 8 (cal 2) | int | 1 | SteRE |  |  |  |  | 1 : in this historical, the calibration of this coef was realized
1142 | 0x0476 | Flash | r | Status of calibration of coefficient 9 (cal 2) | int | 1
1143 | 0x0477 | Flash | r | Status of calibration of coefficient 10 (cal 2) | int | 1
1144 | 0x0478 | Flash | r | Status of calibration of coefficient 11 (cal 2) | int | 1
1145 | 0x0479 | Flash | r | Status of calibration of coefficient 12 (cal 2) | int | 1
1146 | 0x047A | Flash | r | Status of calibration of coefficient 13 (cal 2) | int | 1
1147 | 0x047B | Flash | r | Status of calibration of coefficient 14 (cal 2) | int | 1
1148 | 0x047C | Flash | r | Status of calibration of coefficient 15 (cal 2) | int | 1
1149 | 0x047D | Flash | r | Status of calibration of coefficient 16 (cal 2) | int | 1
1150
1150 | 0x047E | Flash | r | Name of the operator which calibrated the temperature (cal 2) | ascii | 8
1158 | 0x0486 | Flash | r | Date of the calibration of the temperature (cal 2) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
1166 | 0x048E | Flash | r | Name of the operator which calibrated the parameter 1 (cal 2) | ascii | 8
1174 | 0x0496 | Flash | r | Date of the calibration of the parameter 1 (cal 2) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
1182 | 0x049E | Flash | r | Name of the operator which calibrated the parameter 2 (cal 2) | ascii | 8
1190 | 0x04A6 | Flash | r | Date of the calibration of the parameter 2 (cal 2) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
1198 | 0x04AE | Flash | r | Name of the operator which calibrated the parameter 3 (cal 2) | ascii | 8
1206 | 0x04B6 | Flash | r | Date of the calibration of the parameter 3 (cal 2) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
1214 | 0x04BE | Flash | r | Name of the operator which calibrated the parameter 4 (cal 2) | ascii | 8
1222 | 0x04C6 | Flash | r | Date of the calibration of the parameter 4 (cal 2) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
1230 | 0x04CE | Flash | r | Name of the site (cal 2) | ascii | 4 | only using for the turbidity sensor in mg/L and the SS sensor
1234 |  |  |  | Memory location not used |  | 46
1280 | 0x0500 | Flash | r | Value of the temperature standard 1 (cal 3) | float | 2
1282 | 0x0502 | Flash | r | Value of the temperature standard 2 (cal 3) | float | 2
1284 | 0x0504 | Flash | r | Value of standard 1 (cal 3) | float | 2
1286 | 0x0506 | Flash | r | Value of standard 2 (cal 3) | float | 2
1288 | 0x0508 | Flash | r | Value of standard 3 (cal 3) | float | 2
1290 | 0x050A | Flash | r | Value of standard 4 (cal 3) | float | 2
1292 | 0x050C | Flash | r | Value of standard 5 (cal 3) | float | 2
1294 | 0x050E | Flash | r | Value of standard 6 (cal 3) | float | 2
1296 | 0x0510 | Flash | r | Value of standard 7 (cal 3) | float | 2
1298 | 0x0512 | Flash | r | Value of standard 8 (cal 3) | float | 2
1300 | 0x0514 | Flash | r | Value of standard 9 (cal 3) | float | 2
1302 | 0x0516 | Flash | r | Value of standard 10 (cal 3) | float | 2
1304 | 0x0518 | Flash | r | Value of standard 11 (cal 3) | float | 2
1306 | 0x051A | Flash | r | Value of standard 12 (cal 3) | float | 2
1308 | 0x051C | Flash | r | Value of standard 13 (cal 3) | float | 2
1310 | 0x051E | Flash | r | Value of standard 14 (cal 3) | float | 2
1312 | 0x0520 | Flash | r | Value of standard 15 (cal 3) | float | 2
1314 | 0x0522 | Flash | r | Value of standard 16 (cal 3) | float | 2
1316
1316 | 0x0524 | Flash | r | Gross value Offset of the temp (cal 3) | float | 2
1318 | 0x0526 | Flash | r | Gross value gain of the temp (cal 3) | float | 2
1320 | 0x0528 | Flash | r | Gross value 1 (cal 3) | float | 2
1322 | 0x052A | Flash | r | Gross value 2 (cal 3) | float | 2
1324 | 0x052C | Flash | r | Gross value 3 (cal 3) | float | 2
1326 | 0x052E | Flash | r | Gross value 4 (cal 3) | float | 2
1328 | 0x0530 | Flash | r | Gross value 5 (cal 3) | float | 2
1330 | 0x0532 | Flash | r | Gross value 6 (cal 3) | float | 2
1332 | 0x0534 | Flash | r | Gross value 7 (cal 3) | float | 2
1334 | 0x0536 | Flash | r | Gross value 8 (cal 3) | float | 2
1336 | 0x0538 | Flash | r | Gross value 9 (cal 3) | float | 2
1338 | 0x053A | Flash | r | Gross value 10 (cal 3) | float | 2
1340 | 0x053C | Flash | r | Gross value 11 (cal 3) | float | 2
1342 | 0x053E | Flash | r | Gross value 12 (cal 3) | float | 2
1344 | 0x0540 | Flash | r | Gross value 13 (cal 3) | float | 2
1346 | 0x0542 | Flash | r | Gross value 14 (cal 3) | float | 2
1348 | 0x0544 | Flash | r | Gross value 15 (cal 3) | float | 2
1350 | 0x0546 | Flash | r | Gross value 16 (cal 3) | float | 2
1352
1352 | 0x0548 | Flash | r | Offset of the temp (cal 3) | float | 2
1354 | 0x054A | Flash | r | Gain of the temp (cal 3) | float | 2
1356 | 0x054C | Flash | r | Coefficient of correction 1 (cal 3) | float | 2
1358 | 0x054E | Flash | r | Coefficient of correction 2 (cal 3) | float | 2
1360 | 0x0550 | Flash | r | Coefficient of correction 3 (cal 3) | float | 2
1362 | 0x0552 | Flash | r | Coefficient of correction 4 (cal 3) | float | 2
1364 | 0x0554 | Flash | r | Coefficient of correction 5 (cal 3) | float | 2
1366 | 0x0556 | Flash | r | Coefficient of correction 6 (cal 3) | float | 2
1368 | 0x0558 | Flash | r | Coefficient of correction 7 (cal 3) | float | 2
1370 | 0x055A | Flash | r | Coefficient of correction 8 (cal 3) | float | 2
1372 | 0x055C | Flash | r | Coefficient of correction 9 (cal 3) | float | 2
1374 | 0x055E | Flash | r | Coefficient of correction 10 (cal 3) | float | 2
1376 | 0x0560 | Flash | r | Coefficient of correction 11 (cal 3) | float | 2
1378 | 0x0562 | Flash | r | Coefficient of correction 12 (cal 3) | float | 2
1380 | 0x0564 | Flash | r | Coefficient of correction 13 (cal 3) | float | 2
1382 | 0x0566 | Flash | r | Coefficient of correction 14 (cal 3) | float | 2
1384 | 0x0568 | Flash | r | Coefficient of correction 15 (cal 3) | float | 2
1386 | 0x056A | Flash | r | Coefficient of correction 16 (cal 3) | float | 2
1388
1388 | 0x056C | Flash | r | Status of the calibration of the temperature offset (cal 3) | int | 1 | [7 empty] | SteRE |  | SteME | SteHIST | SteCU | SteCH | SteCB | SteE
1389 | 0x056D | Flash | r | Status of the calibration of the temperature gain (cal 3) | int | 1 | SteE : State of the calibration point |  |  |  |  | 00 : calibration point OK (the calibration is not out of range)
1390 | 0x056E | Flash | r | Status of calibration of coefficient 1 (cal 3) | int | 1 | [5 empty] | 01 : calibration point out of range
1391 | 0x056F | Flash | r | Status of calibration of coefficient 2 (cal 3) | int | 1
1392 | 0x0570 | Flash | r | Status of calibration of coefficient 3 (cal 3) | int | 1 | SteCB |  |  |  |  | 1: coef out of range low
1393 | 0x0571 | Flash | r | Status of calibration of coefficient 4 (cal 3) | int | 1 | SteCH |  |  |  |  | 1: coef out of range height
1394 | 0x0572 | Flash | r | Status of calibration of coefficient 5 (cal 3) | int | 1 | SteCU |  |  |  |  | 1 : using of factory coef
1395 | 0x0573 | Flash | r | Status of calibration of coefficient 6 (cal 3) | int | 1 | SteHIST |  |  |  |  | 1 : using of historical coef
1396 | 0x0574 | Flash | r | Status of calibration of coefficient 7 (cal 3) | int | 1 | SteME |  |  |  |  | 1 : recorded measure (differed process)
1397 | 0x0575 | Flash | r | Status of calibration of coefficient 8 (cal 3) | int | 1 | SteRE |  |  |  |  | 1 : in this historical, the calibration of this coef was realized
1398 | 0x0576 | Flash | r | Status of calibration of coefficient 9 (cal 3) | int | 1
1399 | 0x0577 | Flash | r | Status of calibration of coefficient 10 (cal 3) | int | 1
1400 | 0x0578 | Flash | r | Status of calibration of coefficient 11 (cal 3) | int | 1
1401 | 0x0579 | Flash | r | Status of calibration of coefficient 12 (cal 3) | int | 1
1402 | 0x057A | Flash | r | Status of calibration of coefficient 13 (cal 3) | int | 1
1403 | 0x057B | Flash | r | Status of calibration of coefficient 14 (cal 3) | int | 1
1404 | 0x057C | Flash | r | Status of calibration of coefficient 15 (cal 3) | int | 1
1405 | 0x057D | Flash | r | Status of calibration of coefficient 16 (cal 3) | int | 1
1406
1406 | 0x057E | Flash | r | Name of the operator which calibrated the temperature (cal 3) | ascii | 8
1414 | 0x0586 | Flash | r | Date of the calibration of the temperature (cal 3) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
1422 | 0x058E | Flash | r | Name of the operator which calibrated the parameter 1 (cal 3) | ascii | 8
1430 | 0x0596 | Flash | r | Date of the calibration of the parameter 1 (cal 3) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
1438 | 0x059E | Flash | r | Name of the operator which calibrated the parameter 2 (cal 3) | ascii | 8
1446 | 0x05A6 | Flash | r | Date of the calibration of the parameter 2 (cal 3) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
1454 | 0x05AE | Flash | r | Name of the operator which calibrated the parameter 3 (cal 3) | ascii | 8
1462 | 0x05B6 | Flash | r | Date of the calibration of the parameter 3 (cal 3) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
1470 | 0x05BE | Flash | r | Name of the operator which calibrated the parameter 4 (cal 3) | ascii | 8
1478 | 0x05C6 | Flash | r | Date of the calibration of the parameter 4 (cal 3) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
1486 | 0x05CE | Flash | r | Name of the site (cal 3) | ascii | 4 | only using for the turbidity sensor in mg/L and the SS sensor
1490 |  |  |  | Memory location not used |  | 46
1536 | 0x0600 | Flash | r | Value of the temperature standard 1 (cal 4) | float | 2
1538 | 0x0602 | Flash | r | Value of the temperature standard 2 (cal 4) | float | 2
1540 | 0x0604 | Flash | r | Value of standard 1 (cal 4) | float | 2
1542 | 0x0606 | Flash | r | Value of standard 2 (cal 4) | float | 2
1544 | 0x0608 | Flash | r | Value of standard 3 (cal 4) | float | 2
1546 | 0x060A | Flash | r | Value of standard 4 (cal 4) | float | 2
1548 | 0x060C | Flash | r | Value of standard 5 (cal 4) | float | 2
1550 | 0x060E | Flash | r | Value of standard 6 (cal 4) | float | 2
1552 | 0x0610 | Flash | r | Value of standard 7 (cal 4) | float | 2
1554 | 0x0612 | Flash | r | Value of standard 8 (cal 4) | float | 2
1556 | 0x0614 | Flash | r | Value of standard 9 (cal 4) | float | 2
1558 | 0x0616 | Flash | r | Value of standard 10 (cal 4) | float | 2
1560 | 0x0618 | Flash | r | Value of standard 11 (cal 4) | float | 2
1562 | 0x061A | Flash | r | Value of standard 12 (cal 4) | float | 2
1564 | 0x061C | Flash | r | Value of standard 13 (cal 4) | float | 2
1566 | 0x061E | Flash | r | Value of standard 14 (cal 4) | float | 2
1568 | 0x0620 | Flash | r | Value of standard 15 (cal 4) | float | 2
1570 | 0x0622 | Flash | r | Value of standard 16 (cal 4) | float | 2
1572
1572 | 0x0624 | Flash | r | Gross value Offset of the temp (cal 4) | float | 2
1574 | 0x0626 | Flash | r | Gross value gain of the temp (cal 4) | float | 2
1576 | 0x0628 | Flash | r | Gross value 1 (cal 4) | float | 2
1578 | 0x062A | Flash | r | Gross value 2 (cal 4) | float | 2
1580 | 0x062C | Flash | r | Gross value 3 (cal 4) | float | 2
1582 | 0x062E | Flash | r | Gross value 4 (cal 4) | float | 2
1584 | 0x0630 | Flash | r | Gross value 5 (cal 4) | float | 2
1586 | 0x0632 | Flash | r | Gross value 6 (cal 4) | float | 2
1588 | 0x0634 | Flash | r | Gross value 7 (cal 4) | float | 2
1590 | 0x0636 | Flash | r | Gross value 8 (cal 4) | float | 2
1592 | 0x0638 | Flash | r | Gross value 9 (cal 4) | float | 2
1594 | 0x063A | Flash | r | Gross value 10 (cal 4) | float | 2
1596 | 0x063C | Flash | r | Gross value 11 (cal 4) | float | 2
1598 | 0x063E | Flash | r | Gross value 12 (cal 4) | float | 2
1600 | 0x0640 | Flash | r | Gross value 13 (cal 4) | float | 2
1602 | 0x0642 | Flash | r | Gross value 14 (cal 4) | float | 2
1604 | 0x0644 | Flash | r | Gross value 15 (cal 4) | float | 2
1606 | 0x0646 | Flash | r | Gross value 16 (cal 4) | float | 2
1608
1608 | 0x0648 | Flash | r | Offset of the temp (cal 4) | float | 2
1610 | 0x064A | Flash | r | Gain of the temp (cal 4) | float | 2
1612 | 0x064C | Flash | r | Coefficient of correction 1 (cal 4) | float | 2
1614 | 0x064E | Flash | r | Coefficient of correction 2 (cal 4) | float | 2
1616 | 0x0650 | Flash | r | Coefficient of correction 3 (cal 4) | float | 2
1618 | 0x0652 | Flash | r | Coefficient of correction 4 (cal 4) | float | 2
1620 | 0x0654 | Flash | r | Coefficient of correction 5 (cal 4) | float | 2
1622 | 0x0656 | Flash | r | Coefficient of correction 6 (cal 4) | float | 2
1624 | 0x0658 | Flash | r | Coefficient of correction 7 (cal 4) | float | 2
1626 | 0x065A | Flash | r | Coefficient of correction 8 (cal 4) | float | 2
1628 | 0x065C | Flash | r | Coefficient of correction 9 (cal 4) | float | 2
1630 | 0x065E | Flash | r | Coefficient of correction 10 (cal 4) | float | 2
1632 | 0x0660 | Flash | r | Coefficient of correction 11 (cal 4) | float | 2
1634 | 0x0662 | Flash | r | Coefficient of correction 12 (cal 4) | float | 2
1636 | 0x0664 | Flash | r | Coefficient of correction 13 (cal 4) | float | 2
1638 | 0x0666 | Flash | r | Coefficient of correction 14 (cal 4) | float | 2
1640 | 0x0668 | Flash | r | Coefficient of correction 15 (cal 4) | float | 2
1642 | 0x066A | Flash | r | Coefficient of correction 16 (cal 4) | float | 2
1644
1644 | 0x066C | Flash | r | Status of the calibration of the temperature offset (cal 4) | int | 1 | [7 empty] | SteRE |  | SteME | SteHIST | SteCU | SteCH | SteCB | SteE
1645 | 0x066D | Flash | r | Status of the calibration of the temperature gain (cal 4) | int | 1 | SteE : State of the calibration point |  |  |  |  | 00 : calibration point OK (the calibration is not out of range)
1646 | 0x066E | Flash | r | Status of calibration of coefficient 1 (cal 4) | int | 1 | [5 empty] | 01 : calibration point out of range
1647 | 0x066F | Flash | r | Status of calibration of coefficient 2 (cal 4) | int | 1
1648 | 0x0670 | Flash | r | Status of calibration of coefficient 3 (cal 4) | int | 1 | SteCB |  |  |  |  | 1: coef out of range low
1649 | 0x0671 | Flash | r | Status of calibration of coefficient 4 (cal 4) | int | 1 | SteCH |  |  |  |  | 1: coef out of range height
1650 | 0x0672 | Flash | r | Status of calibration of coefficient 5 (cal 4) | int | 1 | SteCU |  |  |  |  | 1 : using of factory coef
1651 | 0x0673 | Flash | r | Status of calibration of coefficient 6 (cal 4) | int | 1 | SteHIST |  |  |  |  | 1 : using of historical coef
1652 | 0x0674 | Flash | r | Status of calibration of coefficient 7 (cal 4) | int | 1 | SteME |  |  |  |  | 1 : recorded measure (differed process)
1653 | 0x0675 | Flash | r | Status of calibration of coefficient 8 (cal 4) | int | 1 | SteRE |  |  |  |  | 1 : in this historical, the calibration of this coef was realized
1654 | 0x0676 | Flash | r | Status of calibration of coefficient 9 (cal 4) | int | 1
1655 | 0x0677 | Flash | r | Status of calibration of coefficient 10 (cal 4) | int | 1
1656 | 0x0678 | Flash | r | Status of calibration of coefficient 11 (cal 4) | int | 1
1657 | 0x0679 | Flash | r | Status of calibration of coefficient 12 (cal 4) | int | 1
1658 | 0x067A | Flash | r | Status of calibration of coefficient 13 (cal 4) | int | 1
1659 | 0x067B | Flash | r | Status of calibration of coefficient 14 (cal 4) | int | 1
1660 | 0x067C | Flash | r | Status of calibration of coefficient 15 (cal 4) | int | 1
1661 | 0x067D | Flash | r | Status of calibration of coefficient 16 (cal 4) | int | 1
1662
1662 | 0x067E | Flash | r | Name of the operator which calibrated the temperature (cal 4) | ascii | 8
1670 | 0x0686 | Flash | r | Date of the calibration of the temperature (cal 4) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
1678 | 0x068E | Flash | r | Name of the operator which calibrated the parameter 1 (cal 4) | ascii | 8
1686 | 0x0696 | Flash | r | Date of the calibration of the parameter 1 (cal 4) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
1694 | 0x069E | Flash | r | Name of the operator which calibrated the parameter 2 (cal 4) | ascii | 8
1702 | 0x06A6 | Flash | r | Date of the calibration of the parameter 2 (cal 4) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
1710 | 0x06AE | Flash | r | Name of the operator which calibrated the parameter 3 (cal 4) | ascii | 8
1718 | 0x06B6 | Flash | r | Date of the calibration of the parameter 3 (cal 4) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
1726 | 0x06BE | Flash | r | Name of the operator which calibrated the parameter 4 (cal 4) | ascii | 8
1734 | 0x06C6 | Flash | r | Date of the calibration of the parameter 4 (cal 4) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
1742 | 0x06CE | Flash | r | Name of the site (cal 4) | ascii | 4 | only using for the turbidity sensor in mg/L and the SS sensor
1746 |  |  |  | Memory location not used |  | 46
1792 | 0x0700 | Flash | r | Value of the temperature standard 1 (cal 5) | float | 2
1794 | 0x0702 | Flash | r | Value of the temperature standard 2 (cal 5) | float | 2
1796 | 0x0704 | Flash | r | Value of standard 1 (cal 5) | float | 2
1798 | 0x0706 | Flash | r | Value of standard 2 (cal 5) | float | 2
1800 | 0x0708 | Flash | r | Value of standard 3 (cal 5) | float | 2
1802 | 0x070A | Flash | r | Value of standard 4 (cal 5) | float | 2
1804 | 0x070C | Flash | r | Value of standard 5 (cal 5) | float | 2
1806 | 0x070E | Flash | r | Value of standard 6 (cal 5) | float | 2
1808 | 0x0710 | Flash | r | Value of standard 7 (cal 5) | float | 2
1810 | 0x0712 | Flash | r | Value of standard 8 (cal 5) | float | 2
1812 | 0x0714 | Flash | r | Value of standard 9 (cal 5) | float | 2
1814 | 0x0716 | Flash | r | Value of standard 10 (cal 5) | float | 2
1816 | 0x0718 | Flash | r | Value of standard 11 (cal 5) | float | 2
1818 | 0x071A | Flash | r | Value of standard 12 (cal 5) | float | 2
1820 | 0x071C | Flash | r | Value of standard 13 (cal 5) | float | 2
1822 | 0x071E | Flash | r | Value of standard 14 (cal 5) | float | 2
1824 | 0x0720 | Flash | r | Value of standard 15 (cal 5) | float | 2
1826 | 0x0722 | Flash | r | Value of standard 16 (cal 5) | float | 2
1828
1828 | 0x0724 | Flash | r | Gross value Offset of the temp (cal 5) | float | 2
1830 | 0x0726 | Flash | r | Gross value gain of the temp (cal 5) | float | 2
1832 | 0x0728 | Flash | r | Gross value 1 (cal 5) | float | 2
1834 | 0x072A | Flash | r | Gross value 2 (cal 5) | float | 2
1836 | 0x072C | Flash | r | Gross value 3 (cal 5) | float | 2
1838 | 0x072E | Flash | r | Gross value 4 (cal 5) | float | 2
1840 | 0x0730 | Flash | r | Gross value 5 (cal 5) | float | 2
1842 | 0x0732 | Flash | r | Gross value 6 (cal 5) | float | 2
1844 | 0x0734 | Flash | r | Gross value 7 (cal 5) | float | 2
1846 | 0x0736 | Flash | r | Gross value 8 (cal 5) | float | 2
1848 | 0x0738 | Flash | r | Gross value 9 (cal 5) | float | 2
1850 | 0x073A | Flash | r | Gross value 10 (cal 5) | float | 2
1852 | 0x073C | Flash | r | Gross value 11 (cal 5) | float | 2
1854 | 0x073E | Flash | r | Gross value 12 (cal 5) | float | 2
1856 | 0x0740 | Flash | r | Gross value 13 (cal 5) | float | 2
1858 | 0x0742 | Flash | r | Gross value 14 (cal 5) | float | 2
1860 | 0x0744 | Flash | r | Gross value 15 (cal 5) | float | 2
1862 | 0x0746 | Flash | r | Gross value 16 (cal 5) | float | 2
1864
1864 | 0x0748 | Flash | r | Offset of the temp (cal 5) | float | 2
1866 | 0x074A | Flash | r | Gain of the temp (cal 5) | float | 2
1868 | 0x074C | Flash | r | Coefficient of correction 1 (cal 5) | float | 2
1870 | 0x074E | Flash | r | Coefficient of correction 2 (cal 5) | float | 2
1872 | 0x0750 | Flash | r | Coefficient of correction 3 (cal 5) | float | 2
1874 | 0x0752 | Flash | r | Coefficient of correction 4 (cal 5) | float | 2
1876 | 0x0754 | Flash | r | Coefficient of correction 5 (cal 5) | float | 2
1878 | 0x0756 | Flash | r | Coefficient of correction 6 (cal 5) | float | 2
1880 | 0x0758 | Flash | r | Coefficient of correction 7 (cal 5) | float | 2
1882 | 0x075A | Flash | r | Coefficient of correction 8 (cal 5) | float | 2
1884 | 0x075C | Flash | r | Coefficient of correction 9 (cal 5) | float | 2
1886 | 0x075E | Flash | r | Coefficient of correction 10 (cal 5) | float | 2
1888 | 0x0760 | Flash | r | Coefficient of correction 11 (cal 5) | float | 2
1890 | 0x0762 | Flash | r | Coefficient of correction 12 (cal 5) | float | 2
1892 | 0x0764 | Flash | r | Coefficient of correction 13 (cal 5) | float | 2
1894 | 0x0766 | Flash | r | Coefficient of correction 14 (cal 5) | float | 2
1896 | 0x0768 | Flash | r | Coefficient of correction 15 (cal 5) | float | 2
1898 | 0x076A | Flash | r | Coefficient of correction 16 (cal 5) | float | 2
1900
1900 | 0x076C | Flash | r | Status of the calibration of the temperature offset (cal 5) | int | 1
1901 | 0x076D | Flash | r | Status of the calibration of the temperature gain (cal 5) | int | 1 | [7 empty] | SteRE |  | SteME | SteHIST | SteCU | SteCH | SteCB | SteE
1902 | 0x076E | Flash | r | Status of calibration of coefficient 1 (cal 5) | int | 1 | SteE : State of the calibration point |  |  |  |  | 00 : calibration point OK (the calibration is not out of range)
1903 | 0x076F | Flash | r | Status of calibration of coefficient 2 (cal 5) | int | 1 | [5 empty] | 01 : calibration point out of range
1904 | 0x0770 | Flash | r | Status of calibration of coefficient 3 (cal 5) | int | 1
1905 | 0x0771 | Flash | r | Status of calibration of coefficient 4 (cal 5) | int | 1 | SteCB |  |  |  |  | 1: coef out of range low
1906 | 0x0772 | Flash | r | Status of calibration of coefficient 5 (cal 5) | int | 1 | SteCH |  |  |  |  | 1: coef out of range height
1907 | 0x0773 | Flash | r | Status of calibration of coefficient 6 (cal 5) | int | 1 | SteCU |  |  |  |  | 1 : using of factory coef
1908 | 0x0774 | Flash | r | Status of calibration of coefficient 7 (cal 5) | int | 1 | SteHIST |  |  |  |  | 1 : using of historical coef
1909 | 0x0775 | Flash | r | Status of calibration of coefficient 8 (cal 5) | int | 1 | SteME |  |  |  |  | 1 : recorded measure (differed process)
1910 | 0x0776 | Flash | r | Status of calibration of coefficient 9 (cal 5) | int | 1 | SteRE |  |  |  |  | 1 : in this historical, the calibration of this coef was realized
1911 | 0x0777 | Flash | r | Status of calibration of coefficient 10 (cal 5) | int | 1
1912 | 0x0778 | Flash | r | Status of calibration of coefficient 11 (cal 5) | int | 1
1913 | 0x0779 | Flash | r | Status of calibration of coefficient 12 (cal 5) | int | 1
1914 | 0x077A | Flash | r | Status of calibration of coefficient 13 (cal 5) | int | 1
1915 | 0x077B | Flash | r | Status of calibration of coefficient 14 (cal 5) | int | 1
1916 | 0x077C | Flash | r | Status of calibration of coefficient 15 (cal 5) | int | 1
1917 | 0x077D | Flash | r | Status of calibration of coefficient 16 (cal 5) | int | 1
1918
1918 | 0x077E | Flash | r | Name of the operator which calibrated the temperature (cal 5) | ascii | 8
1926 | 0x0786 | Flash | r | Date of the calibration of the temperature (cal 5) | ascii | 8
1934 | 0x078E | Flash | r | Name of the operator which calibrated the parameter 1 (cal 5) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
1942 | 0x0796 | Flash | r | Date of the calibration of the parameter 1 (cal 5) | ascii | 8
1950 | 0x079E | Flash | r | Name of the operator which calibrated the parameter 2 (cal 5) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
1958 | 0x07A6 | Flash | r | Date of the calibration of the parameter 2 (cal 5) | ascii | 8
1966 | 0x07AE | Flash | r | Name of the operator which calibrated the parameter 3 (cal 5) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
1974 | 0x07B6 | Flash | r | Date of the calibration of the parameter 3 (cal 5) | ascii | 8
1982 | 0x07BE | Flash | r | Name of the operator which calibrated the parameter 4 (cal 5) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
1990 | 0x07C6 | Flash | r | Date of the calibration of the parameter 4 (cal 5) | ascii | 8
1998 | 0x07CE | Flash | r | Name of the site (cal 5) | ascii | 4 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
2002 |  |  |  | Memory location not used |  | 46 | only using for the turbidity sensor in mg/L and the SS sensor
2048 | 0x0800 | Flash | r | Value of the temperature standard 1 (cal 6) | float | 2
2050 | 0x0802 | Flash | r | Value of the temperature standard 2 (cal 6) | float | 2
2052 | 0x0804 | Flash | r | Value of standard 1 (cal 6) | float | 2
2054 | 0x0806 | Flash | r | Value of standard 2 (cal 6) | float | 2
2056 | 0x0808 | Flash | r | Value of standard 3 (cal 6) | float | 2
2058 | 0x080A | Flash | r | Value of standard 4 (cal 6) | float | 2
2060 | 0x080C | Flash | r | Value of standard 5 (cal 6) | float | 2
2062 | 0x080E | Flash | r | Value of standard 6 (cal 6) | float | 2
2064 | 0x0810 | Flash | r | Value of standard 7 (cal 6) | float | 2
2066 | 0x0812 | Flash | r | Value of standard 8 (cal 6) | float | 2
2068 | 0x0814 | Flash | r | Value of standard 9 (cal 6) | float | 2
2070 | 0x0816 | Flash | r | Value of standard 10 (cal 6) | float | 2
2072 | 0x0818 | Flash | r | Value of standard 11 (cal 6) | float | 2
2074 | 0x081A | Flash | r | Value of standard 12 (cal 6) | float | 2
2076 | 0x081C | Flash | r | Value of standard 13 (cal 6) | float | 2
2078 | 0x081E | Flash | r | Value of standard 14 (cal 6) | float | 2
2080 | 0x0820 | Flash | r | Value of standard 15 (cal 6) | float | 2
2082 | 0x0822 | Flash | r | Value of standard 16 (cal 6) | float | 2
2084
2084 | 0x0824 | Flash | r | Gross value Offset of the temp (cal 6) | float | 2
2086 | 0x0826 | Flash | r | Gross value gain of the temp (cal 6) | float | 2
2088 | 0x0828 | Flash | r | Gross value 1 (cal 6) | float | 2
2090 | 0x082A | Flash | r | Gross value 2 (cal 6) | float | 2
2092 | 0x082C | Flash | r | Gross value 3 (cal 6) | float | 2
2094 | 0x082E | Flash | r | Gross value 4 (cal 6) | float | 2
2096 | 0x0830 | Flash | r | Gross value 5 (cal 6) | float | 2
2098 | 0x0832 | Flash | r | Gross value 6 (cal 6) | float | 2
2100 | 0x0834 | Flash | r | Gross value 7 (cal 6) | float | 2
2102 | 0x0836 | Flash | r | Gross value 8 (cal 6) | float | 2
2104 | 0x0838 | Flash | r | Gross value 9 (cal 6) | float | 2
2106 | 0x083A | Flash | r | Gross value 10 (cal 6) | float | 2
2108 | 0x083C | Flash | r | Gross value 11 (cal 6) | float | 2
2110 | 0x083E | Flash | r | Gross value 12 (cal 6) | float | 2
2112 | 0x0840 | Flash | r | Gross value 13 (cal 6) | float | 2
2114 | 0x0842 | Flash | r | Gross value 14 (cal 6) | float | 2
2116 | 0x0844 | Flash | r | Gross value 15 (cal 6) | float | 2
2118 | 0x0846 | Flash | r | Gross value 16 (cal 6) | float | 2
2120
2120 | 0x0848 | Flash | r | Offset of the temp (cal 6) | float | 2
2122 | 0x084A | Flash | r | Gain of the temp (cal 6) | float | 2
2124 | 0x084C | Flash | r | Coefficient of correction 1 (cal 6) | float | 2
2126 | 0x084E | Flash | r | Coefficient of correction 2 (cal 6) | float | 2
2128 | 0x0850 | Flash | r | Coefficient of correction 3 (cal 6) | float | 2
2130 | 0x0852 | Flash | r | Coefficient of correction 4 (cal 6) | float | 2
2132 | 0x0854 | Flash | r | Coefficient of correction 5 (cal 6) | float | 2
2134 | 0x0856 | Flash | r | Coefficient of correction 6 (cal 6) | float | 2
2136 | 0x0858 | Flash | r | Coefficient of correction 7 (cal 6) | float | 2
2138 | 0x085A | Flash | r | Coefficient of correction 8 (cal 6) | float | 2
2140 | 0x085C | Flash | r | Coefficient of correction 9 (cal 6) | float | 2
2142 | 0x085E | Flash | r | Coefficient of correction 10 (cal 6) | float | 2
2144 | 0x0860 | Flash | r | Coefficient of correction 11 (cal 6) | float | 2
2146 | 0x0862 | Flash | r | Coefficient of correction 12 (cal 6) | float | 2
2148 | 0x0864 | Flash | r | Coefficient of correction 13 (cal 6) | float | 2
2150 | 0x0866 | Flash | r | Coefficient of correction 14 (cal 6) | float | 2
2152 | 0x0868 | Flash | r | Coefficient of correction 15 (cal 6) | float | 2
2154 | 0x086A | Flash | r | Coefficient of correction 16 (cal 6) | float | 2
2156
2156 | 0x086C | Flash | r | Status of the calibration of the temperature offset (cal 6) | int | 1 | [7 empty] | SteRE |  | SteME | SteHIST | SteCU | SteCH | SteCB | SteE
2157 | 0x086D | Flash | r | Status of the calibration of the temperature gain (cal 6) | int | 1 | SteE : State of the calibration point |  |  |  |  | 00 : calibration point OK (the calibration is not out of range)
2158 | 0x086E | Flash | r | Status of calibration of coefficient 1 (cal 6) | int | 1 | [5 empty] | 01 : calibration point out of range
2159 | 0x086F | Flash | r | Status of calibration of coefficient 2 (cal 6) | int | 1
2160 | 0x0870 | Flash | r | Status of calibration of coefficient 3 (cal 6) | int | 1 | SteCB |  |  |  |  | 1: coef out of range low
2161 | 0x0871 | Flash | r | Status of calibration of coefficient 4 (cal 6) | int | 1 | SteCH |  |  |  |  | 1: coef out of range height
2162 | 0x0872 | Flash | r | Status of calibration of coefficient 5 (cal 6) | int | 1 | SteCU |  |  |  |  | 1 : using of factory coef
2163 | 0x0873 | Flash | r | Status of calibration of coefficient 6 (cal 6) | int | 1 | SteHIST |  |  |  |  | 1 : using of historical coef
2164 | 0x0874 | Flash | r | Status of calibration of coefficient 7 (cal 6) | int | 1 | SteME |  |  |  |  | 1 : recorded measure (differed process)
2165 | 0x0875 | Flash | r | Status of calibration of coefficient 8 (cal 6) | int | 1 | SteRE |  |  |  |  | 1 : in this historical, the calibration of this coef was realized
2166 | 0x0876 | Flash | r | Status of calibration of coefficient 9 (cal 6) | int | 1
2167 | 0x0877 | Flash | r | Status of calibration of coefficient 10 (cal 6) | int | 1
2168 | 0x0878 | Flash | r | Status of calibration of coefficient 11 (cal 6) | int | 1
2169 | 0x0879 | Flash | r | Status of calibration of coefficient 12 (cal 6) | int | 1
2170 | 0x087A | Flash | r | Status of calibration of coefficient 13 (cal 6) | int | 1
2171 | 0x087B | Flash | r | Status of calibration of coefficient 14 (cal 6) | int | 1
2172 | 0x087C | Flash | r | Status of calibration of coefficient 15 (cal 6) | int | 1
2173 | 0x087D | Flash | r | Status of calibration of coefficient 16 (cal 6) | int | 1
2174
2174 | 0x087E | Flash | r | Name of the operator which calibrated the temperature (cal 6) | ascii | 8
2182 | 0x0886 | Flash | r | Date of the calibration of the temperature (cal 6) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
2190 | 0x088E | Flash | r | Name of the operator which calibrated the parameter 1 (cal 6) | ascii | 8
2198 | 0x0896 | Flash | r | Date of the calibration of the parameter 1 (cal 6) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
2206 | 0x089E | Flash | r | Name of the operator which calibrated the parameter 2 (cal 6) | ascii | 8
2214 | 0x08A6 | Flash | r | Date of the calibration of the parameter 2 (cal 6) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
2222 | 0x08AE | Flash | r | Name of the operator which calibrated the parameter 3 (cal 6) | ascii | 8
2230 | 0x08B6 | Flash | r | Date of the calibration of the parameter 3 (cal 6) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
2238 | 0x08BE | Flash | r | Name of the operator which calibrated the parameter 4 (cal 6) | ascii | 8
2246 | 0x08C6 | Flash | r | Date of the calibration of the parameter 4 (cal 6) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
2254 | 0x08CE | Flash | r | Name of the site (cal 6) | ascii | 4 | only using for the turbidity sensor in mg/L and the SS sensor
2258 |  |  |  | Memory location not used |  | 46
2304 | 0x0900 | Flash | r | Value of the temperature standard 1 (cal 7) | float | 2
2306 | 0x0902 | Flash | r | Value of the temperature standard 2 (cal 7) | float | 2
2308 | 0x0904 | Flash | r | Value of standard 1 (cal 7) | float | 2
2310 | 0x0906 | Flash | r | Value of standard 2 (cal 7) | float | 2
2312 | 0x0908 | Flash | r | Value of standard 3 (cal 7) | float | 2
2314 | 0x090A | Flash | r | Value of standard 4 (cal 7) | float | 2
2316 | 0x090C | Flash | r | Value of standard 5 (cal 7) | float | 2
2318 | 0x090E | Flash | r | Value of standard 6 (cal 7) | float | 2
2320 | 0x0910 | Flash | r | Value of standard 7 (cal 7) | float | 2
2322 | 0x0912 | Flash | r | Value of standard 8 (cal 7) | float | 2
2324 | 0x0914 | Flash | r | Value of standard 9 (cal 7) | float | 2
2326 | 0x0916 | Flash | r | Value of standard 10 (cal 7) | float | 2
2328 | 0x0918 | Flash | r | Value of standard 11 (cal 7) | float | 2
2330 | 0x091A | Flash | r | Value of standard 12 (cal 7) | float | 2
2332 | 0x091C | Flash | r | Value of standard 13 (cal 7) | float | 2
2334 | 0x091E | Flash | r | Value of standard 14 (cal 7) | float | 2
2336 | 0x0920 | Flash | r | Value of standard 15 (cal 7) | float | 2
2338 | 0x0922 | Flash | r | Value of standard 16 (cal 7) | float | 2
2340
2340 | 0x0924 | Flash | r | Gross value Offset of the temp (cal 7) | float | 2
2342 | 0x0926 | Flash | r | Gross value gain of the temp (cal 7) | float | 2
2344 | 0x0928 | Flash | r | Gross value 1 (cal 7) | float | 2
2346 | 0x092A | Flash | r | Gross value 2 (cal 7) | float | 2
2348 | 0x092C | Flash | r | Gross value 3 (cal 7) | float | 2
2350 | 0x092E | Flash | r | Gross value 4 (cal 7) | float | 2
2352 | 0x0930 | Flash | r | Gross value 5 (cal 7) | float | 2
2354 | 0x0932 | Flash | r | Gross value 6 (cal 7) | float | 2
2356 | 0x0934 | Flash | r | Gross value 7 (cal 7) | float | 2
2358 | 0x0936 | Flash | r | Gross value 8 (cal 7) | float | 2
2360 | 0x0938 | Flash | r | Gross value 9 (cal 7) | float | 2
2362 | 0x093A | Flash | r | Gross value 10 (cal 7) | float | 2
2364 | 0x093C | Flash | r | Gross value 11 (cal 7) | float | 2
2366 | 0x093E | Flash | r | Gross value 12 (cal 7) | float | 2
2368 | 0x0940 | Flash | r | Gross value 13 (cal 7) | float | 2
2370 | 0x0942 | Flash | r | Gross value 14 (cal 7) | float | 2
2372 | 0x0944 | Flash | r | Gross value 15 (cal 7) | float | 2
2374 | 0x0946 | Flash | r | Gross value 16 (cal 7) | float | 2
2376
2376 | 0x0948 | Flash | r | Offset of the temp (cal 7) | float | 2
2378 | 0x094A | Flash | r | Gain of the temp (cal 7) | float | 2
2380 | 0x094C | Flash | r | Coefficient of correction 1 (cal 7) | float | 2
2382 | 0x094E | Flash | r | Coefficient of correction 2 (cal 7) | float | 2
2384 | 0x0950 | Flash | r | Coefficient of correction 3 (cal 7) | float | 2
2386 | 0x0952 | Flash | r | Coefficient of correction 4 (cal 7) | float | 2
2388 | 0x0954 | Flash | r | Coefficient of correction 5 (cal 7) | float | 2
2390 | 0x0956 | Flash | r | Coefficient of correction 6 (cal 7) | float | 2
2392 | 0x0958 | Flash | r | Coefficient of correction 7 (cal 7) | float | 2
2394 | 0x095A | Flash | r | Coefficient of correction 8 (cal 7) | float | 2
2396 | 0x095C | Flash | r | Coefficient of correction 9 (cal 7) | float | 2
2398 | 0x095E | Flash | r | Coefficient of correction 10 (cal 7) | float | 2
2400 | 0x0960 | Flash | r | Coefficient of correction 11 (cal 7) | float | 2
2402 | 0x0962 | Flash | r | Coefficient of correction 12 (cal 7) | float | 2
2404 | 0x0964 | Flash | r | Coefficient of correction 13 (cal 7) | float | 2
2406 | 0x0966 | Flash | r | Coefficient of correction 14 (cal 7) | float | 2
2408 | 0x0968 | Flash | r | Coefficient of correction 15 (cal 7) | float | 2
2410 | 0x096A | Flash | r | Coefficient of correction 16 (cal 7) | float | 2
2412
2412 | 0x096C | Flash | r | Status of the calibration of the temperature offset (cal 7) | int | 1 | [7 empty] | SteRE |  | SteME | SteHIST | SteCU | SteCH | SteCB | SteE
2413 | 0x096D | Flash | r | Status of the calibration of the temperature gain (cal 7) | int | 1 | SteE : State of the calibration point |  |  |  |  | 00 : calibration point OK (the calibration is not out of range)
2414 | 0x096E | Flash | r | Status of calibration of coefficient 1 (cal 7) | int | 1 | [5 empty] | 01 : calibration point out of range
2415 | 0x096F | Flash | r | Status of calibration of coefficient 2 (cal 7) | int | 1
2416 | 0x0970 | Flash | r | Status of calibration of coefficient 3 (cal 7) | int | 1 | SteCB |  |  |  |  | 1: coef out of range low
2417 | 0x0971 | Flash | r | Status of calibration of coefficient 4 (cal 7) | int | 1 | SteCH |  |  |  |  | 1: coef out of range height
2418 | 0x0972 | Flash | r | Status of calibration of coefficient 5 (cal 7) | int | 1 | SteCU |  |  |  |  | 1 : using of factory coef
2419 | 0x0973 | Flash | r | Status of calibration of coefficient 6 (cal 7) | int | 1 | SteHIST |  |  |  |  | 1 : using of historical coef
2420 | 0x0974 | Flash | r | Status of calibration of coefficient 7 (cal 7) | int | 1 | SteME |  |  |  |  | 1 : recorded measure (differed process)
2421 | 0x0975 | Flash | r | Status of calibration of coefficient 8 (cal 7) | int | 1 | SteRE |  |  |  |  | 1 : in this historical, the calibration of this coef was realized
2422 | 0x0976 | Flash | r | Status of calibration of coefficient 9 (cal 7) | int | 1
2423 | 0x0977 | Flash | r | Status of calibration of coefficient 10 (cal 7) | int | 1
2424 | 0x0978 | Flash | r | Status of calibration of coefficient 11 (cal 7) | int | 1
2425 | 0x0979 | Flash | r | Status of calibration of coefficient 12 (cal 7) | int | 1
2426 | 0x097A | Flash | r | Status of calibration of coefficient 13 (cal 7) | int | 1
2427 | 0x097B | Flash | r | Status of calibration of coefficient 14 (cal 7) | int | 1
2428 | 0x097C | Flash | r | Status of calibration of coefficient 15 (cal 7) | int | 1
2429 | 0x097D | Flash | r | Status of calibration of coefficient 16 (cal 7) | int | 1
2430
2430 | 0x097E | Flash | r | Name of the operator which calibrated the temperature (cal 7) | ascii | 8
2438 | 0x0986 | Flash | r | Date of the calibration of the temperature (cal 7) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
2446 | 0x098E | Flash | r | Name of the operator which calibrated the parameter 1 (cal 7) | ascii | 8
2454 | 0x0996 | Flash | r | Date of the calibration of the parameter 1 (cal 7) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
2462 | 0x099E | Flash | r | Name of the operator which calibrated the parameter 2 (cal 7) | ascii | 8
2470 | 0x09A6 | Flash | r | Date of the calibration of the parameter 2 (cal 7) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
2478 | 0x09AE | Flash | r | Name of the operator which calibrated the parameter 3 (cal 7) | ascii | 8
2486 | 0x09B6 | Flash | r | Date of the calibration of the parameter 3 (cal 7) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
2494 | 0x09BE | Flash | r | Name of the operator which calibrated the parameter 4 (cal 7) | ascii | 8
2502 | 0x09C6 | Flash | r | Date of the calibration of the parameter 4 (cal 7) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
2510 | 0x09CE | Flash | r | Name of the site (cal 7) | ascii | 4 | only using for the turbidity sensor in mg/L and the SS sensor
2514 |  |  |  | Memory location not used |  | 46
2560 | 0x0A00 | Flash | r | Value of the temperature standard 1 (cal 8) | float | 2
2562 | 0x0A02 | Flash | r | Value of the temperature standard 2 (cal 8) | float | 2
2564 | 0x0A04 | Flash | r | Value of standard 1 (cal 8) | float | 2
2566 | 0x0A06 | Flash | r | Value of standard 2 (cal 8) | float | 2
2568 | 0x0A08 | Flash | r | Value of standard 3 (cal 8) | float | 2
2570 | 0x0A0A | Flash | r | Value of standard 4 (cal 8) | float | 2
2572 | 0x0A0C | Flash | r | Value of standard 5 (cal 8) | float | 2
2574 | 0x0A0E | Flash | r | Value of standard 6 (cal 8) | float | 2
2576 | 0x0A10 | Flash | r | Value of standard 7 (cal 8) | float | 2
2578 | 0x0A12 | Flash | r | Value of standard 8 (cal 8) | float | 2
2580 | 0x0A14 | Flash | r | Value of standard 9 (cal 8) | float | 2
2582 | 0x0A16 | Flash | r | Value of standard 10 (cal 8) | float | 2
2584 | 0x0A18 | Flash | r | Value of standard 11 (cal 8) | float | 2
2586 | 0x0A1A | Flash | r | Value of standard 12 (cal 8) | float | 2
2588 | 0x0A1C | Flash | r | Value of standard 13 (cal 8) | float | 2
2590 | 0x0A1E | Flash | r | Value of standard 14 (cal 8) | float | 2
2592 | 0x0A20 | Flash | r | Value of standard 15 (cal 8) | float | 2
2594 | 0x0A22 | Flash | r | Value of standard 16 (cal 8) | float | 2
2596
2596 | 0x0A24 | Flash | r | Gross value Offset of the temp (cal 8) | float | 2
2598 | 0x0A26 | Flash | r | Gross value gain of the temp (cal 8) | float | 2
2600 | 0x0A28 | Flash | r | Gross value 1 (cal 8) | float | 2
2602 | 0x0A2A | Flash | r | Gross value 2 (cal 8) | float | 2
2604 | 0x0A2C | Flash | r | Gross value 3 (cal 8) | float | 2
2606 | 0x0A2E | Flash | r | Gross value 4 (cal 8) | float | 2
2608 | 0x0A30 | Flash | r | Gross value 5 (cal 8) | float | 2
2610 | 0x0A32 | Flash | r | Gross value 6 (cal 8) | float | 2
2612 | 0x0A34 | Flash | r | Gross value 7 (cal 8) | float | 2
2614 | 0x0A36 | Flash | r | Gross value 8 (cal 8) | float | 2
2616 | 0x0A38 | Flash | r | Gross value 9 (cal 8) | float | 2
2618 | 0x0A3A | Flash | r | Gross value 10 (cal 8) | float | 2
2620 | 0x0A3C | Flash | r | Gross value 11 (cal 8) | float | 2
2622 | 0x0A3E | Flash | r | Gross value 12 (cal 8) | float | 2
2624 | 0x0A40 | Flash | r | Gross value 13 (cal 8) | float | 2
2626 | 0x0A42 | Flash | r | Gross value 14 (cal 8) | float | 2
2628 | 0x0A44 | Flash | r | Gross value 15 (cal 8) | float | 2
2630 | 0x0A46 | Flash | r | Gross value 16 (cal 8) | float | 2
2632
2632 | 0x0A48 | Flash | r | Offset of the temp (cal 8) | float | 2
2634 | 0x0A4A | Flash | r | Gain of the temp (cal 8) | float | 2
2636 | 0x0A4C | Flash | r | Coefficient of correction 1 (cal 8) | float | 2
2638 | 0x0A4E | Flash | r | Coefficient of correction 2 (cal 8) | float | 2
2640 | 0x0A50 | Flash | r | Coefficient of correction 3 (cal 8) | float | 2
2642 | 0x0A52 | Flash | r | Coefficient of correction 4 (cal 8) | float | 2
2644 | 0x0A54 | Flash | r | Coefficient of correction 5 (cal 8) | float | 2
2646 | 0x0A56 | Flash | r | Coefficient of correction 6 (cal 8) | float | 2
2648 | 0x0A58 | Flash | r | Coefficient of correction 7 (cal 8) | float | 2
2650 | 0x0A5A | Flash | r | Coefficient of correction 8 (cal 8) | float | 2
2652 | 0x0A5C | Flash | r | Coefficient of correction 9 (cal 8) | float | 2
2654 | 0x0A5E | Flash | r | Coefficient of correction 10 (cal 8) | float | 2
2656 | 0x0A60 | Flash | r | Coefficient of correction 11 (cal 8) | float | 2
2658 | 0x0A62 | Flash | r | Coefficient of correction 12 (cal 8) | float | 2
2660 | 0x0A64 | Flash | r | Coefficient of correction 13 (cal 8) | float | 2
2662 | 0x0A66 | Flash | r | Coefficient of correction 14 (cal 8) | float | 2
2664 | 0x0A68 | Flash | r | Coefficient of correction 15 (cal 8) | float | 2
2666 | 0x0A6A | Flash | r | Coefficient of correction 16 (cal 8) | float | 2
2668
2668 | 0x0A6C | Flash | r | Status of the calibration of the temperature offset (cal 8) | int | 1 | [7 empty] | SteRE |  | SteME | SteHIST | SteCU | SteCH | SteCB | SteE
2669 | 0x0A6D | Flash | r | Status of the calibration of the temperature gain (cal 8) | int | 1 | SteE : State of the calibration point |  |  |  |  | 00 : calibration point OK (the calibration is not out of range)
2670 | 0x0A6E | Flash | r | Status of calibration of coefficient 1 (cal 8) | int | 1 | [5 empty] | 01 : calibration point out of range
2671 | 0x0A6F | Flash | r | Status of calibration of coefficient 2 (cal 8) | int | 1
2672 | 0x0A70 | Flash | r | Status of calibration of coefficient 3 (cal 8) | int | 1 | SteCB |  |  |  |  | 1: coef out of range low
2673 | 0x0A71 | Flash | r | Status of calibration of coefficient 4 (cal 8) | int | 1 | SteCH |  |  |  |  | 1: coef out of range height
2674 | 0x0A72 | Flash | r | Status of calibration of coefficient 5 (cal 8) | int | 1 | SteCU |  |  |  |  | 1 : using of factory coef
2675 | 0x0A73 | Flash | r | Status of calibration of coefficient 6 (cal 8) | int | 1 | SteHIST |  |  |  |  | 1 : using of historical coef
2676 | 0x0A74 | Flash | r | Status of calibration of coefficient 7 (cal 8) | int | 1 | SteME |  |  |  |  | 1 : recorded measure (differed process)
2677 | 0x0A75 | Flash | r | Status of calibration of coefficient 8 (cal 8) | int | 1 | SteRE |  |  |  |  | 1 : in this historical, the calibration of this coef was realized
2678 | 0x0A76 | Flash | r | Status of calibration of coefficient 9 (cal 8) | int | 1
2679 | 0x0A77 | Flash | r | Status of calibration of coefficient 10 (cal 8) | int | 1
2680 | 0x0A78 | Flash | r | Status of calibration of coefficient 11 (cal 8) | int | 1
2681 | 0x0A79 | Flash | r | Status of calibration of coefficient 12 (cal 8) | int | 1
2682 | 0x0A7A | Flash | r | Status of calibration of coefficient 13 (cal 8) | int | 1
2683 | 0x0A7B | Flash | r | Status of calibration of coefficient 14 (cal 8) | int | 1
2684 | 0x0A7C | Flash | r | Status of calibration of coefficient 15 (cal 8) | int | 1
2685 | 0x0A7D | Flash | r | Status of calibration of coefficient 16 (cal 8) | int | 1
2686
2686 | 0x0A7E | Flash | r | Name of the operator which calibrated the temperature (cal 8) | ascii | 8
2694 | 0x0A86 | Flash | r | Date of the calibration of the temperature (cal 8) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
2702 | 0x0A8E | Flash | r | Name of the operator which calibrated the parameter 1 (cal 8) | ascii | 8
2710 | 0x0A96 | Flash | r | Date of the calibration of the parameter 1 (cal 8) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
2718 | 0x0A9E | Flash | r | Name of the operator which calibrated the parameter 2 (cal 8) | ascii | 8
2726 | 0x0AA6 | Flash | r | Date of the calibration of the parameter 2 (cal 8) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
2734 | 0x0AAE | Flash | r | Name of the operator which calibrated the parameter 3 (cal 8) | ascii | 8
2742 | 0x0AB6 | Flash | r | Date of the calibration of the parameter 3 (cal 8) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
2750 | 0x0ABE | Flash | r | Name of the operator which calibrated the parameter 4 (cal 8) | ascii | 8
2758 | 0x0AC6 | Flash | r | Date of the calibration of the parameter 4 (cal 8) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
2766 | 0x0ACE | Flash | r | Name of the site (cal 8) | ascii | 4 | only using for the turbidity sensor in mg/L and the SS sensor
2770 |  |  |  | Memory location not used |  | 46
2816 | 0x0B00 | Flash | r | Value of the temperature standard 1 (cal 9) | float | 2
2818 | 0x0B02 | Flash | r | Value of the temperature standard 2 (cal 9) | float | 2
2820 | 0x0B04 | Flash | r | Value of standard 1 (cal 9) | float | 2
2822 | 0x0B06 | Flash | r | Value of standard 2 (cal 9) | float | 2
2824 | 0x0B08 | Flash | r | Value of standard 3 (cal 9) | float | 2
2826 | 0x0B0A | Flash | r | Value of standard 4 (cal 9) | float | 2
2828 | 0x0B0C | Flash | r | Value of standard 5 (cal 9) | float | 2
2830 | 0x0B0E | Flash | r | Value of standard 6 (cal 9) | float | 2
2832 | 0x0B10 | Flash | r | Value of standard 7 (cal 9) | float | 2
2834 | 0x0B12 | Flash | r | Value of standard 8 (cal 9) | float | 2
2836 | 0x0B14 | Flash | r | Value of standard 9 (cal 9) | float | 2
2838 | 0x0B16 | Flash | r | Value of standard 10 (cal 9) | float | 2
2840 | 0x0B18 | Flash | r | Value of standard 11 (cal 9) | float | 2
2842 | 0x0B1A | Flash | r | Value of standard 12 (cal 9) | float | 2
2844 | 0x0B1C | Flash | r | Value of standard 13 (cal 9) | float | 2
2846 | 0x0B1E | Flash | r | Value of standard 14 (cal 9) | float | 2
2848 | 0x0B20 | Flash | r | Value of standard 15 (cal 9) | float | 2
2850 | 0x0B22 | Flash | r | Value of standard 16 (cal 9) | float | 2
2852
2852 | 0x0B24 | Flash | r | Gross value Offset of the temp (cal 9) | float | 2
2854 | 0x0B26 | Flash | r | Gross value gain of the temp (cal 9) | float | 2
2856 | 0x0B28 | Flash | r | Gross value 1 (cal 9) | float | 2
2858 | 0x0B2A | Flash | r | Gross value 2 (cal 9) | float | 2
2860 | 0x0B2C | Flash | r | Gross value 3 (cal 9) | float | 2
2862 | 0x0B2E | Flash | r | Gross value 4 (cal 9) | float | 2
2864 | 0x0B30 | Flash | r | Gross value 5 (cal 9) | float | 2
2866 | 0x0B32 | Flash | r | Gross value 6 (cal 9) | float | 2
2868 | 0x0B34 | Flash | r | Gross value 7 (cal 9) | float | 2
2870 | 0x0B36 | Flash | r | Gross value 8 (cal 9) | float | 2
2872 | 0x0B38 | Flash | r | Gross value 9 (cal 9) | float | 2
2874 | 0x0B3A | Flash | r | Gross value 10 (cal 9) | float | 2
2876 | 0x0B3C | Flash | r | Gross value 11 (cal 9) | float | 2
2878 | 0x0B3E | Flash | r | Gross value 12 (cal 9) | float | 2
2880 | 0x0B40 | Flash | r | Gross value 13 (cal 9) | float | 2
2882 | 0x0B42 | Flash | r | Gross value 14 (cal 9) | float | 2
2884 | 0x0B44 | Flash | r | Gross value 15 (cal 9) | float | 2
2886 | 0x0B46 | Flash | r | Gross value 16 (cal 9) | float | 2
2888
2888 | 0x0B48 | Flash | r | Offset of the temp (cal 9) | float | 2
2890 | 0x0B4A | Flash | r | Gain of the temp (cal 9) | float | 2
2892 | 0x0B4C | Flash | r | Coefficient of correction 1 (cal 9) | float | 2
2894 | 0x0B4E | Flash | r | Coefficient of correction 2 (cal 9) | float | 2
2896 | 0x0B50 | Flash | r | Coefficient of correction 3 (cal 9) | float | 2
2898 | 0x0B52 | Flash | r | Coefficient of correction 4 (cal 9) | float | 2
2900 | 0x0B54 | Flash | r | Coefficient of correction 5 (cal 9) | float | 2
2902 | 0x0B56 | Flash | r | Coefficient of correction 6 (cal 9) | float | 2
2904 | 0x0B58 | Flash | r | Coefficient of correction 7 (cal 9) | float | 2
2906 | 0x0B5A | Flash | r | Coefficient of correction 8 (cal 9) | float | 2
2908 | 0x0B5C | Flash | r | Coefficient of correction 9 (cal 9) | float | 2
2910 | 0x0B5E | Flash | r | Coefficient of correction 10 (cal 9) | float | 2
2912 | 0x0B60 | Flash | r | Coefficient of correction 11 (cal 9) | float | 2
2914 | 0x0B62 | Flash | r | Coefficient of correction 12 (cal 9) | float | 2
2916 | 0x0B64 | Flash | r | Coefficient of correction 13 (cal 9) | float | 2
2918 | 0x0B66 | Flash | r | Coefficient of correction 14 (cal 9) | float | 2
2920 | 0x0B68 | Flash | r | Coefficient of correction 15 (cal 9) | float | 2
2922 | 0x0B6A | Flash | r | Coefficient of correction 16 (cal 9) | float | 2
2924
2924 | 0x0B6C | Flash | r | Status of the calibration of the temperature offset (cal 9) | int | 1 | [7 empty] | SteRE |  | SteME | SteHIST | SteCU | SteCH | SteCB | SteE
2925 | 0x0B6D | Flash | r | Status of the calibration of the temperature gain (cal 9) | int | 1 | SteE : State of the calibration point |  |  |  |  | 00 : calibration point OK (the calibration is not out of range)
2926 | 0x0B6E | Flash | r | Status of calibration of coefficient 1 (cal 9) | int | 1 | [5 empty] | 01 : calibration point out of range
2927 | 0x0B6F | Flash | r | Status of calibration of coefficient 2 (cal 9) | int | 1
2928 | 0x0B70 | Flash | r | Status of calibration of coefficient 3 (cal 9) | int | 1 | SteCB |  |  |  |  | 1: coef out of range low
2929 | 0x0B71 | Flash | r | Status of calibration of coefficient 4 (cal 9) | int | 1 | SteCH |  |  |  |  | 1: coef out of range height
2930 | 0x0B72 | Flash | r | Status of calibration of coefficient 5 (cal 9) | int | 1 | SteCU |  |  |  |  | 1 : using of factory coef
2931 | 0x0B73 | Flash | r | Status of calibration of coefficient 6 (cal 9) | int | 1 | SteHIST |  |  |  |  | 1 : using of historical coef
2932 | 0x0B74 | Flash | r | Status of calibration of coefficient 7 (cal 9) | int | 1 | SteME |  |  |  |  | 1 : recorded measure (differed process)
2933 | 0x0B75 | Flash | r | Status of calibration of coefficient 8 (cal 9) | int | 1 | SteRE |  |  |  |  | 1 : in this historical, the calibration of this coef was realized
2934 | 0x0B76 | Flash | r | Status of calibration of coefficient 9 (cal 9) | int | 1
2935 | 0x0B77 | Flash | r | Status of calibration of coefficient 10 (cal 9) | int | 1
2936 | 0x0B78 | Flash | r | Status of calibration of coefficient 11 (cal 9) | int | 1
2937 | 0x0B79 | Flash | r | Status of calibration of coefficient 12 (cal 9) | int | 1
2938 | 0x0B7A | Flash | r | Status of calibration of coefficient 13 (cal 9) | int | 1
2939 | 0x0B7B | Flash | r | Status of calibration of coefficient 14 (cal 9) | int | 1
2940 | 0x0B7C | Flash | r | Status of calibration of coefficient 15 (cal 9) | int | 1
2941 | 0x0B7D | Flash | r | Status of calibration of coefficient 16 (cal 9) | int | 1
2942
2942 | 0x0B7E | Flash | r | Name of the operator which calibrated the temperature (cal 9) | ascii | 8
2950 | 0x0B86 | Flash | r | Date of the calibration of the temperature (cal 9) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
2958 | 0x0B8E | Flash | r | Name of the operator which calibrated the parameter 1 (cal 9) | ascii | 8
2966 | 0x0B96 | Flash | r | Date of the calibration of the parameter 1 (cal 9) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
2974 | 0x0B9E | Flash | r | Name of the operator which calibrated the parameter 2 (cal 9) | ascii | 8
2982 | 0x0BA6 | Flash | r | Date of the calibration of the parameter 2 (cal 9) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
2990 | 0x0BAE | Flash | r | Name of the operator which calibrated the parameter 3 (cal 9) | ascii | 8
2998 | 0x0BB6 | Flash | r | Date of the calibration of the parameter 3 (cal 9) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
3006 | 0x0BBE | Flash | r | Name of the operator which calibrated the parameter 4 (cal 9) | ascii | 8
3014 | 0x0BC6 | Flash | r | Date of the calibration of the parameter 4 (cal 9) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
3022 | 0x0BCE | Flash | r | Name of the site (cal 9) | ascii | 4 | only using for the turbidity sensor in mg/L and the SS sensor
3026 |  |  |  | Memory location not used |  | 46
3072 | 0x0C00 | Flash | r | Value of the temperature standard 1 (cal 10) | float | 2
3074 | 0x0C02 | Flash | r | Value of the temperature standard 2 (cal 10) | float | 2
3076 | 0x0C04 | Flash | r | Value of standard 1 (cal 10) | float | 2
3078 | 0x0C06 | Flash | r | Value of standard 2 (cal 10) | float | 2
3080 | 0x0C08 | Flash | r | Value of standard 3 (cal 10) | float | 2
3082 | 0x0C0A | Flash | r | Value of standard 4 (cal 10) | float | 2
3084 | 0x0C0C | Flash | r | Value of standard 5 (cal 10) | float | 2
3086 | 0x0C0E | Flash | r | Value of standard 6 (cal 10) | float | 2
3088 | 0x0C10 | Flash | r | Value of standard 7 (cal 10) | float | 2
3090 | 0x0C12 | Flash | r | Value of standard 8 (cal 10) | float | 2
3092 | 0x0C14 | Flash | r | Value of standard 9 (cal 10) | float | 2
3094 | 0x0C16 | Flash | r | Value of standard 10 (cal 10) | float | 2
3096 | 0x0C18 | Flash | r | Value of standard 11 (cal 10) | float | 2
3098 | 0x0C1A | Flash | r | Value of standard 12 (cal 10) | float | 2
3100 | 0x0C1C | Flash | r | Value of standard 13 (cal 10) | float | 2
3102 | 0x0C1E | Flash | r | Value of standard 14 (cal 10) | float | 2
3104 | 0x0C20 | Flash | r | Value of standard 15 (cal 10) | float | 2
3106 | 0x0C22 | Flash | r | Value of standard 16 (cal 10) | float | 2
3108
3108 | 0x0C24 | Flash | r | Gross value Offset of the temp (cal 10) | float | 2
3110 | 0x0C26 | Flash | r | Gross value gain of the temp (cal 10) | float | 2
3112 | 0x0C28 | Flash | r | Gross value 1 (cal 10) | float | 2
3114 | 0x0C2A | Flash | r | Gross value 2 (cal 10) | float | 2
3116 | 0x0C2C | Flash | r | Gross value 3 (cal 10) | float | 2
3118 | 0x0C2E | Flash | r | Gross value 4 (cal 10) | float | 2
3120 | 0x0C30 | Flash | r | Gross value 5 (cal 10) | float | 2
3122 | 0x0C32 | Flash | r | Gross value 6 (cal 10) | float | 2
3124 | 0x0C34 | Flash | r | Gross value 7 (cal 10) | float | 2
3126 | 0x0C36 | Flash | r | Gross value 8 (cal 10) | float | 2
3128 | 0x0C38 | Flash | r | Gross value 9 (cal 10) | float | 2
3130 | 0x0C3A | Flash | r | Gross value 10 (cal 10) | float | 2
3132 | 0x0C3C | Flash | r | Gross value 11 (cal 10) | float | 2
3134 | 0x0C3E | Flash | r | Gross value 12 (cal 10) | float | 2
3136 | 0x0C40 | Flash | r | Gross value 13 (cal 10) | float | 2
3138 | 0x0C42 | Flash | r | Gross value 14 (cal 10) | float | 2
3140 | 0x0C44 | Flash | r | Gross value 15 (cal 10) | float | 2
3142 | 0x0C46 | Flash | r | Gross value 16 (cal 10) | float | 2
3144
3144 | 0x0C48 | Flash | r | Offset of the temp (cal 10) | float | 2
3146 | 0x0C4A | Flash | r | Gain of the temp (cal 10) | float | 2
3148 | 0x0C4C | Flash | r | Coefficient of correction 1 (cal 10) | float | 2
3150 | 0x0C4E | Flash | r | Coefficient of correction 2 (cal 10) | float | 2
3152 | 0x0C50 | Flash | r | Coefficient of correction 3 (cal 10) | float | 2
3154 | 0x0C52 | Flash | r | Coefficient of correction 4 (cal 10) | float | 2
3156 | 0x0C54 | Flash | r | Coefficient of correction 5 (cal 10) | float | 2
3158 | 0x0C56 | Flash | r | Coefficient of correction 6 (cal 10) | float | 2
3160 | 0x0C58 | Flash | r | Coefficient of correction 7 (cal 10) | float | 2
3162 | 0x0C5A | Flash | r | Coefficient of correction 8 (cal 10) | float | 2
3164 | 0x0C5C | Flash | r | Coefficient of correction 9 (cal 10) | float | 2
3166 | 0x0C5E | Flash | r | Coefficient of correction 10 (cal 10) | float | 2
3168 | 0x0C60 | Flash | r | Coefficient of correction 11 (cal 10) | float | 2
3170 | 0x0C62 | Flash | r | Coefficient of correction 12 (cal 10) | float | 2
3172 | 0x0C64 | Flash | r | Coefficient of correction 13 (cal 10) | float | 2
3174 | 0x0C66 | Flash | r | Coefficient of correction 14 (cal 10) | float | 2
3176 | 0x0C68 | Flash | r | Coefficient of correction 15 (cal 10) | float | 2
3178 | 0x0C6A | Flash | r | Coefficient of correction 16 (cal 10) | float | 2
3180
3180 | 0x0C6C | Flash | r | Status of the calibration of the temperature offset (cal 10) | int | 1 | [7 empty] | SteRE |  | SteME | SteHIST | SteCU | SteCH | SteCB | SteE
3181 | 0x0C6D | Flash | r | Status of the calibration of the temperature gain (cal 10) | int | 1 | SteE : State of the calibration point |  |  |  |  | 00 : calibration point OK (the calibration is not out of range)
3182 | 0x0C6E | Flash | r | Status of calibration of coefficient 1 (cal 10) | int | 1 | [5 empty] | 01 : calibration point out of range
3183 | 0x0C6F | Flash | r | Status of calibration of coefficient 2 (cal 10) | int | 1
3184 | 0x0C70 | Flash | r | Status of calibration of coefficient 3 (cal 10) | int | 1 | SteCB |  |  |  |  | 1: coef out of range low
3185 | 0x0C71 | Flash | r | Status of calibration of coefficient 4 (cal 10) | int | 1 | SteCH |  |  |  |  | 1: coef out of range height
3186 | 0x0C72 | Flash | r | Status of calibration of coefficient 5 (cal 10) | int | 1 | SteCU |  |  |  |  | 1 : using of factory coef
3187 | 0x0C73 | Flash | r | Status of calibration of coefficient 6 (cal 10) | int | 1 | SteHIST |  |  |  |  | 1 : using of historical coef
3188 | 0x0C74 | Flash | r | Status of calibration of coefficient 7 (cal 10) | int | 1 | SteME |  |  |  |  | 1 : recorded measure (differed process)
3189 | 0x0C75 | Flash | r | Status of calibration of coefficient 8 (cal 10) | int | 1 | SteRE |  |  |  |  | 1 : in this historical, the calibration of this coef was realized
3190 | 0x0C76 | Flash | r | Status of calibration of coefficient 9 (cal 10) | int | 1
3191 | 0x0C77 | Flash | r | Status of calibration of coefficient 10 (cal 10) | int | 1
3192 | 0x0C78 | Flash | r | Status of calibration of coefficient 11 (cal 10) | int | 1
3193 | 0x0C79 | Flash | r | Status of calibration of coefficient 12 (cal 10) | int | 1
3194 | 0x0C7A | Flash | r | Status of calibration of coefficient 13 (cal 10) | int | 1
3195 | 0x0C7B | Flash | r | Status of calibration of coefficient 14 (cal 10) | int | 1
3196 | 0x0C7C | Flash | r | Status of calibration of coefficient 15 (cal 10) | int | 1
3197 | 0x0C7D | Flash | r | Status of calibration of coefficient 16 (cal 10) | int | 1
3198
3198 | 0x0C7E | Flash | r | Name of the operator which calibrated the temperature (cal 10) | ascii | 8
3206 | 0x0C86 | Flash | r | Date of the calibration of the temperature (cal 10) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
3214 | 0x0C8E | Flash | r | Name of the operator which calibrated the parameter 1 (cal 10) | ascii | 8
3222 | 0x0C96 | Flash | r | Date of the calibration of the parameter 1 (cal 10) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
3230 | 0x0C9E | Flash | r | Name of the operator which calibrated the parameter 2 (cal 10) | ascii | 8
3238 | 0x0CA6 | Flash | r | Date of the calibration of the parameter 2 (cal 10) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
3246 | 0x0CAE | Flash | r | Name of the operator which calibrated the parameter 3 (cal 10) | ascii | 8
3254 | 0x0CB6 | Flash | r | Date of the calibration of the parameter 3 (cal 10) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
3262 | 0x0CBE | Flash | r | Name of the operator which calibrated the parameter 4 (cal 10) | ascii | 8
3270 | 0x0CC6 | Flash | r | Date of the calibration of the parameter 4 (cal 10) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
3278 | 0x0CCE | Flash | r | Name of the site (cal 10) | ascii | 4 | only using for the turbidity sensor in mg/L and the SS sensor
3282 |  |  |  | Memory location not used |  | 46
3328 | 0x0D00 | ROM | r | POD description | ascii | 16 | (32 characters)
3344 | 0x0D10 | ROM | r | Serial number | ascii | 16
3360 | 0x0D20 | ROM | r | Software version: | int | 1 | Heavy Weight: Major version and Light Weight: Minor Version.
3361 | 0x0D21 | ROM | r | Hardware version: | int | 1 | Heavy Weight: Major version and Light Weight: Minor Version.
3362 | 0x0D22 | ROM | r | Temperature unit | ascii | 3
3365 | 0x0D25 | ROM | r | Parameter 1 unit | ascii | 3
3368 | 0x0D28 | ROM | r | Parameter 2 unit | ascii | 3
3371 | 0x0D2B | ROM | r | Parameter 3 unit | ascii | 3
3374 | 0x0D2E | ROM | r | Parameter 4 unit | ascii | 3
3377 | 0x0D31 | ROM | r | Factory Offset of the temp | float | 2
3379 | 0x0D33 | ROM | r | Factory Gain of the temp | float | 2
3381 | 0x0D35 | ROM | r | Factory Coefficient of correction 1 | float | 2
3383 | 0x0D37 | ROM | r | Factory Coefficient of correction 2 | float | 2
3385 | 0x0D39 | ROM | r | Factory Coefficient of correction 3 | float | 2
3387 | 0x0D3B | ROM | r | Factory Coefficient of correction 4 | float | 2
3389 | 0x0D3D | ROM | r | Factory Coefficient of correction 5 | float | 2
3391 | 0x0D3F | ROM | r | Factory Coefficient of correction 6 | float | 2
3393 | 0x0D41 | ROM | r | Factory Coefficient of correction 7 | float | 2
3395 | 0x0D43 | ROM | r | Factory Coefficient of correction 8 | float | 2
3397 | 0x0D45 | ROM | r | Factory Coefficient of correction 9 | float | 2
3399 | 0x0D47 | ROM | r | Factory Coefficient of correction 10 | float | 2
3401 | 0x0D49 | ROM | r | Factory Coefficient of correction 11 | float | 2
3403 | 0x0D4B | ROM | r | Factory Coefficient of correction 12 | float | 2
3405 | 0x0D4D | ROM | r | Factory Coefficient of correction 13 | float | 2
3407 | 0x0D4F | ROM | r | Factory Coefficient of correction 14 | float | 2
3409 | 0x0D51 | ROM | r | Factory Coefficient of correction 15 | float | 2
3411 | 0x0D53 | ROM | r | Factory Coefficient of correction 16 | float | 2
3413
3413 | 0x0D55 | ROM | r | Name of the operator which calibrated the temperature (factory) | ascii | 8
3421 | 0x0D5D | ROM | r | Date of the calibration of the temperature (factory) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
3429 | 0x0D65 | ROM | r | Name of the operator which calibrated the parameter 1 (factory) | ascii | 8
3437 | 0x0D6D | ROM | r | Date of the calibration of the parameter 1 (factory) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
3445 | 0x0D75 | ROM | r | Name of the operator which calibrated the parameter 2 (factory) | ascii | 8
3453 | 0x0D7D | ROM | r | Date of the calibration of the parameter 2 (factory) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
3461 | 0x0D85 | ROM | r | Name of the operator which calibrated the parameter 3 (factory) | ascii | 8
3469 | 0x0D8D | ROM | r | Date of the calibration of the parameter 3 (factory) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
3477 | 0x0D95 | ROM | r | Name of the operator which calibrated the parameter 4 (factory) | ascii | 8
3485 | 0x0D9D | ROM | r | Date of the calibration of the parameter 4 (factory) | ascii | 8 | Minutes |  | Hour |  | Day |  | Month |  | Year |  |  |  | Space | Space | Space | Space
3493
3493 | 0x0DA5 | ROM | r | Low limit of the temperature offset | float | 2
3495 | 0x0DA7 | ROM | r | High limit of the temperature offset | float | 2
3497 | 0x0DA9 | ROM | r | Low limit of the temperature gain | float | 2
3499 | 0x0DAB | ROM | r | High limit of the temperature gain | float | 2
3501 | 0x0DAD | ROM | r | Low limit of the coefficient of correction 1 | float | 2
3503 | 0x0DAF | ROM | r | High limit of the coefficient of correction 1 | float | 2
3505 | 0x0DB1 | ROM | r | Low limit of the coefficient of correction 2 | float | 2
3507 | 0x0DB3 | ROM | r | High limit of the coefficient of correction 2 | float | 2
3509 | 0x0DB5 | ROM | r | Low limit of the coefficient of correction 3 | float | 2
3511 | 0x0DB7 | ROM | r | High limit of the coefficient of correction 3 | float | 2
3513 | 0x0DB9 | ROM | r | Low limit of the coefficient of correction 4 | float | 2
3515 | 0x0DBB | ROM | r | High limit of the coefficient of correction 4 | float | 2
3517 | 0x0DBD | ROM | r | Low limit of the coefficient of correction 5 | float | 2
3519 | 0x0DBF | ROM | r | High limit of the coefficient of correction 5 | float | 2
3521 | 0x0DC1 | ROM | r | Low limit of the coefficient of correction 6 | float | 2
3523 | 0x0DC3 | ROM | r | High limit of the coefficient of correction 6 | float | 2
3525 | 0x0DC5 | ROM | r | Low limit of the coefficient of correction 7 | float | 2
3527 | 0x0DC7 | ROM | r | High limit of the coefficient of correction 7 | float | 2
3529 | 0x0DC9 | ROM | r | Low limit of the coefficient of correction 8 | float | 2
3531 | 0x0DCB | ROM | r | High limit of the coefficient of correction 8 | float | 2
3533 | 0x0DCD | ROM | r | Low limit of the coefficient of correction 9 | float | 2
3535 | 0x0DCF | ROM | r | High limit of the coefficient of correction 9 | float | 2
3537 | 0x0DD1 | ROM | r | Low limit of the coefficient of correction 10 | float | 2
3539 | 0x0DD3 | ROM | r | High limit of the coefficient of correction 10 | float | 2
3541 | 0x0DD5 | ROM | r | Low limit of the coefficient of correction 11 | float | 2
3543 | 0x0DD7 | ROM | r | High limit of the coefficient of correction 11 | float | 2
3545 | 0x0DD9 | ROM | r | Low limit of the coefficient of correction 12 | float | 2
3547 | 0x0DDB | ROM | r | High limit of the coefficient of correction 12 | float | 2
3549 | 0x0DDD | ROM | r | Low limit of the coefficient of correction 13 | float | 2
3551 | 0x0DDF | ROM | r | High limit of the coefficient of correction 13 | float | 2
3553 | 0x0DE1 | ROM | r | Low limit of the coefficient of correction 14 | float | 2
3555 | 0x0DE3 | ROM | r | High limit of the coefficient of correction 14 | float | 2
3557 | 0x0DE5 | ROM | r | Low limit of the coefficient of correction 15 | float | 2
3559 | 0x0DE7 | ROM | r | High limit of the coefficient of correction 15 | float | 2
3561 | 0x0DE9 | ROM | r | Low limit of the coefficient of correction 16 | float | 2
3563 | 0x0DEB | ROM | r | High limit of the coefficient of correction 16 | float | 2
3565
3565 | 0x0DED | ROM | r | High limit of the mesure of temperature | float | 2
3567 | 0x0DEF | ROM | r | Low limit of the mesure of temperature | float | 2
3569 | 0x0DF1 | ROM | r | Low limit of the measure of the parameter 1 | float | 2
3571 | 0x0DF3 | ROM | r | High limit of the measure of the parameter 1 | float | 2
3573 | 0x0DF5 | ROM | r | Low limit of the measure of the parameter 2 | float | 2
3575 | 0x0DF7 | ROM | r | High limit of the measure of the parameter 2 | float | 2
3577 | 0x0DF9 | ROM | r | Low limit of the measure of the parameter 3 | float | 2
3579 | 0x0DFB | ROM | r | High limit of the measure of the parameter 3 | float | 2
3581 | 0x0DFD | ROM | r | Low limit of the measure of the parameter 4 | float | 2
3583 | 0x0DFF | ROM | r | High limit of the measure of the parameter 4 | float | 2
3585
3585
3585
3585 |  | Flash | r | Date of manufacture of the lot of electronic card | ascii | 5 | Format : Semxx-20xx
3590 |  | Flash | r | Name of the operator which made the functional test of the electronic card | ascii | 4 | 8 digits
3594 |  | Flash | r | Date of the functional test of the electronic card | ascii | 5 | Format : xx/xx/20xx (day / month / year)
3599 |  | Flash | r | Number of electronic card made in the lot | int | 1
3600 |  | Flash | r | Number of the electronic card in the lot of manufacturing | int | 1
3601 |  | Flash | r | Number of file of industrialization + Indication having allowed the manufacturing of the card | ascii | 5 | Format : DIxxxxxx.x
3606 |  | Flash | r | Soft Version of the autotest of the micro | int | 1 | Strong weight: Majeur version and Weak weight: minor Version.
3607 |  | Flash | r | Excel soft Version : Test_fonctionnel_des_pods_Vx.xls | int | 1
3608
3608 |  |  |  | Memory location not used |  | 71
3679
3679
3679
3679 | 0x0E5F | Flash | Except Modbus | utilisé pour enregistrer la derniere gamme utilisé | int | 1
3680 | 0x0E60 | Flash | Except Modbus | Averaging board for the Parameter 1 | float | 104
3784 | 0x0EC8 | Flash | Except Modbus | Averaging board for the Parameter 2 | float | 104
3888 | 0x0F30 | Flash | Except Modbus | Averaging board for the Parameter 3 | float | 104
3992 | 0x0F98 | Flash | Except Modbus | Averaging board for the Parameter 4 | float | 104
4096
4096 | 0x1000
4098 | 0x1002 | Flash | r | Variable 1 | float | 2
4100 | 0x1004 | Flash | r | Variable 2 | float | 2
4102 | 0x1006 | Flash | r | Variable 3 | float | 2
4104 | 0x1008 | Flash | r | Variable 4 | float | 2
4106 | 0x100A | Flash | r | Variable 5 | float | 2
4108 | 0X100C | Flash | r | Variable 6 | float | 2
4110 | 0x100E | Flash | r | Variable 7 | float | 2
4112 | 0x1010 | Flash | r | Variable 8 | float | 2
4114 | 0x1012 | Flash | r | Variable 9 | float | 2
4116 | 0x1014 | Flash | r | Variable 10 | float | 2
4118 | 0x1016 | Flash | r | Variable 11 | float | 2
4120 | 0x1018 | Flash | r | Variable 12 | float | 2
4122 | 0x101A | Flash | r | Variable 13 | float | 2
4124 | 0X101C | Flash | r | Variable 14 | float | 2
4126 | 0x101E | Flash | r | Variable 15 | float | 2
4128 |  | Flash | r | Variable 16 | float | 2
4128 | 0x1020
4256 |  |  |  | Memory location not used |  | 128
4256
4256
4256
4256
4256
4256 | 0x10A0

## OPTOD
| POD description | "ODO / Temperature PONSEL " | [5 empty] | Information available
| default address | 10 | [5 empty] | Used for measurement parameter setting
| serial number | SN-PODOA-xxxx | [5 empty] | Used for measurement:
| Default averaging value | 1 | [5 empty] | Used for POD parameter setting:
Used for parameter calibration
System, used exclusively by Ponsel
[7 empty] | NU | Not used
|  | Temperature measurement | Parameter 1 | Parameter 2 | Parameter 3 | Parameter 4
| Parameter measurement: | Temperature | Oxygen | Oxygen | Oxygen | NU
| Unit: | °C | %Sat | mg/L | ppm | NU
| Managed by history or by site: | historical | historical | historical | historical | NU
| Compensates Temperature returned by Master | These compensations are not used for this measurement. | in use, in °C, Default value (when Pod switched on: 25°C) |  |  | NU
| Compensates 1 returned by Master |  | Atmospheric pressure, in hPa, Default value (when Pod switched on: 1013hPa). |  |  | NU
| Compensates 2 returned by Master |  | Salinity, in ppt, Default value (when Pod switched on: 0 ppt). |  |  | NU
| Compensates 3 returned by Master | NU
Measurement configuration: |  | In use | Used for Parameters 1, 2 and 3. | NU | NU | NU
Configures measurement type. | differed Mode in use: | NO | NO | NU | NU | NU
| corrected measurement (default value): | returns measurement | returns measurement | NU | NU | NU
| corrected measurement 1: | NU
| corrected measurement 2: | NU
| gross measurement: | NU
| Range configuration: CgGamme | NU
Measurement status: |  | In use | In use | In use | In use | NU
Measurement status | Measurement OK with INFO 1 | NU | Measurement disrupted | Measurement disrupted | Measurement disrupted | NU
| Measurement OK with INFO 2 | NU | NU | NU | NU | NU
| Measurement impossible with INFO 3 | NU | Tablet missing | Tablet missing | Tablet missing | NU
| Measurement impossible with INFO 4 | NU | NU | NU | NU | NU
Description of calibration coefficients : all coefficients are based on default coefficient.
|  | Information: | Unit: | Coefficient used for parameter: | Comments:
| Offset Temp Coefficient | Temperature offset | °C | Temperature
| Temp Slope Coefficient | Temperature slope | % | Temperature
| Coef 1 | Offset ('A') | Gross Unit | Parameter 1 | Coefficient can be calibrated by integrator.
| Coef 2 | 'B' |  | Paramètre 1 | Coefficient calculé seulement en USINE
| Coef 3 | 'C' |  | Paramètre 1 | Coefficient calculé seulement en USINE
| Coef 4 | Pente ('D') | Gross Unit | Parameter 1 | Coefficient can be calibrated by integrator.
| Coef 5 | 'E' |  | Parametre 1 | Coefficient calculé seulement en USINE
| Coef 6 | 'F' |  | Parametre 1 | Coefficient calculé seulement en USINE
| Coef 7 | NU |  |  | Gross value for temporary use
| Coef 8 | NU |  |  | Gross value for temporary use
| Coef 9 | NU |  |  | Gross value for temporary use
| Coef 10 | NU
| Coef 11 | NU
| Coef 12 | NU
| Coef 13 | NU
| Coef 14 | NU
| Coef 15 | NU
| Coef 16 | NU
Description of calibration standards :
|  | Master returns: | Unit: | Calibration standard used for parameter: | Comments:
| calibration standard 1 for temperature | Calibration standard value (offset): | °C | Temperature
| calibration standard 2 for temperature | Calibration standard value (slope) | °C | Temperature
| standard 1 | 0x00 |  | Parameters 1, 2 and 3 : Oxygen | Sensor must be in the sulfite at stabilized temperature T1, between 5 and 25 °C.
| standard 2 | 0x00 |  | Parameters 1, 2 and 3 : Oxygen
| standard 3 | Calibration standard value | % | Parameters 1, 2 and 3 : Oxygen
| standard 4 | 0x00 |  | Parameters 1, 2 and 3 : Oxygen | Sensor must indicate 100% (when exposed to air) at stabilized temperature T1, between 5 and 25 °C.
| standard 5 | 0x00 |  | Parameters 1, 2 and 3 : Oxygen
| standard 6 | 0x00 |  | Parameters 1, 2 and 3 : Oxygen
| standard 7 | NU
| standard 8 | NU
| standard 9 | NU
| standard 10 | NU
| standard 11 | NU
| standard 12 | NU
| standard 13 | NU
| standard 14 | NU
| standard 15 | NU
| standard 16 | NU

## Turbi
| POD description | "Nephelo/TU / Temperature PONSEL " | [5 empty] | Information available
| default address | 40 | [5 empty] | Used for measurement parameter setting
| serial number | SN-PNEPA-xxxx | [5 empty] | Used for measurement:
| Default averaging value | 10 | [5 empty] | Used for POD parameter setting:
Used for parameter calibration
| Comments : Master cannot ask for Nephelo measurement simultaneously with TU measurement. By default, Nephelo measurement is enabled. | [6 empty] | System, used exclusively by Ponsel
[7 empty] | NU | Not used
|  | Temperature measurement | Parameter 1 | Parameter 2 | Parameter 3 | Parameter 4
| Parameter measurement: | Temperature | Nephelo turbidity | Nephelo turbidity | TU turbidity | NU
| Unit: | °C | NTU | FNU | mg/L | NU
| Managed by history or by site: | historical | historical | historical | Site | NU
| Compensates Temperature returned by Master | NU | used in °C |  |  | NU
| Compensates 1 returned by Master | NU
| Compensates 2 returned by Master | NU
| Compensates 3 returned by Master | NU
Measurement configuration: |  | In use | Used for Parameters 1 and 2. | NU | Used for Parameter 4. | NU
Measurement type configuration | differed Mode in use: | NO | NO | NU | YES | NU
| corrected measurement (by default): | returns measurement | returns measurement | NU | returns measurement | NU
| corrected measurement 1: | NU
| corrected measurement 2: | NU
| gross measurement: | NU
| Range configuration: CgGamme | NU | Use of 4 ranges for Parameters 1 and 2: - range 1: 0 / 50 NTU (FNU) - range 2: 0 / 200 NTU (FNU) - range 3: 0 / 1000 NTU (FNU) - range 4: 0 / 4000 NTU (FNU) | NU | NU | NU
Measurement status: |  | In use | In use | NU | In use | NU
Measurement status | Measurement OK with INFO 1 | NU | Too much daylight. | NU | Too much daylight. | NU
| Measurement OK with INFO 2 | NU | NU | NU | NU | NU
| Measurement impossible with INFO 3 | NU | Too much daylight. | NU | Too much daylight. | NU
| Measurement impossible with INFO 4 | NU | NU | NU | NU | NU
Description of calibration coefficients : all coefficients are based on default coefficient.
|  | Information: | Unit: | Coefficient used for parameter: | Comments:
| Offset Temp Coefficient | Temperature offset | °C | Temperature
| Temp Slope Coefficient | Temperature slope | % | Temperature
| Coef 1 | Range offset 1 | NTU | Parameters 1 and 2
| Coef 2 | Range slope 1 | % | Parameters 1 and 2
| Coef 3 | Range offset 2 | NTU | Parameters 1 and 2
| Coef 4 | Range slope 2 | % | Parameters 1 and 2
| Coef 5 | Range offset 3 | NTU | Parameters 1 and 2
| Coef 6 | Range slope 3 | % | Parameters 1 and 2
| Coef 7 | Range offset 4 | NTU | Parameters 1 and 2
| Coef 8 | Range slope 4 | % | Parameters 1 and 2
| Coef 9 | TU offset | mg/L | Parameter 3
| Coef 10 | TU slope | % | Parameter 3
| Coef 11 | NU
| Coef 12 | NU
| Coef 13 | NU
| Coef 14 | NU
| Coef 15 | NU
| Coef 16 | NU
Description of calibration standards :
|  | Master returns: | Unit: | Calibration standard used for parameter: | Comments:
| calibration standard 1 for temperature | Calibration standard value (offset): | °C | Temperature | value recommended by Ponsel: near 0°C
| calibration standard 2 for temperature | Calibration standard value (slope) | °C | Temperature | value recommended by Ponsel: near 25°C
| standard 1 | Calibration standard value (offset): | in NTUs | Parameters 1 and 2 | value recommended by Ponsel: demineralised water 0 NTU
| standard 2 | Calibration standard value (slope) | in NTUs | Parameters 1 and 2 | value recommended by Ponsel: 25 NTU (middle of range)
| standard 3 | Calibration standard value (offset): | in NTUs | Parameters 1 and 2 | value recommended by Ponsel: demineralised water 0 NTU
| standard 4 | Calibration standard value (slope) | in NTUs | Parameters 1 and 2 | value recommended by Ponsel: 100 NTU (middle of range)
| standard 5 | Calibration standard value (offset): | in NTUs | Parameters 1 and 2 | value recommended by Ponsel: demineralised water 0 NTU
| standard 6 | Calibration standard value (slope) | in NTUs | Parameters 1 and 2 | value recommended by Ponsel: 500 NTU (middle of range)
| standard 7 | Calibration standard value (offset): | in NTUs | Parameters 1 and 2 | value recommended by Ponsel: demineralised water 0 NTU
| standard 8 | Calibration standard value (slope) | in NTUs | Parameters 1 and 2 | value recommended by Ponsel: 2000 NTU (middle of range)
| standard 9 | Calibration standard value (offset): | in mg/L | Parameter 3 | ??? To be verified
| standard 10 | Calibration standard value (slope) | in mg/L | Parameter 3 | ??? To be verified
| standard 11 | NU
| standard 12 | NU
| standard 13 | NU
| standard 14 | NU
| standard 15 | NU
| standard 16 | NU

## pH_Redox
| POD description | "pH / Redox / Temperature PONSEL " | [5 empty] | Information available
| default address | 20 | [5 empty] | Used for measurement parameter setting
| serial number | SN-PPHRA-xxxx | [5 empty] | Used for measurement:
| Default averaging value | 1 | [5 empty] | Used for POD parameter setting:
Used for parameter calibration
System, used exclusively by Ponsel
[7 empty] | NU | Not used
|  | Temperature measurement | Parameter 1 | Parameter 2 | Parameter 3 | Parameter 4
| Parameter measurement: | Temperature | pH | Redox | NU | NU
| Unit: | °C | in pH units | in mV | NU | NU
| Managed by history or by site: | historical | historical | historical | NU | NU
| Compensates Temperature returned by Master | NU | used in °C | NU | NU | NU
| Compensates 1 returned by Master | NU
| Compensates 2 returned by Master | NU
| Compensates 3 returned by Master | NU
Measurement configuration: |  | In use | In use | In use | NU | NU
Measurement type configuration | differed Mode in use: | NO | NO | NO | NU | NU
| corrected measurement (by default): | returns measurement | returns measurement | returns measurement | NU | NU
| corrected measurement 1: | NU
| corrected measurement 2: | NU
| gross measurement: | NU
| Range configuration: CgGamme | NU
Measurement status: |  | In use | In use | In use | NU | NU
Measurement status | Measurement OK with INFO 1 | NU | NU | NU | NU | NU
| Measurement OK with INFO 2 | NU | NU | NU | NU | NU
| Measurement impossible with INFO 3 | NU | NU | NU | NU | NU
| Measurement impossible with INFO 4 | NU | NU | NU | NU | NU
Description of calibration coefficients : all coefficients are based on default coefficient.
|  | Information: | Unit: | Coefficient used for parameter: | Comments:
| Offset Temp Coefficient | Temperature offset | °C | Temperature
| Temp Slope Coefficient | Temperature slope | % | Temperature
| Coef 1 | pH offset | in pH units | Parameter 1
| Coef 2 | pH slope | % | Parameter 1
| Coef 3 | Redox offset | in mV | Parameter 2
| Coef 4 | Redox slope | % | Parameter 2
| Coef 5 | Used for pH buffer: | pH | Parameter 1
| Coef 6 | NU
| Coef 7 | NU
| Coef 8 | NU
| Coef 9 | NU
| Coef 10 | NU
| Coef 11 | NU
| Coef 12 | NU
| Coef 13 | NU
| Coef 14 | NU
| Coef 15 | NU
| Coef 16 | NU
Description of calibration standards :
|  | Master returns: | Unit: | Calibration standard used for parameter: | Comments:
| calibration standard 1 for temperature | Calibration standard value (offset): | °C | Temperature | value recommended by Ponsel: near 0°C
| calibration standard 2 for temperature | Calibration standard value (slope) | °C | Temperature | value recommended by Ponsel: near 25°C
| standard 1 | Calibration standard value (offset): | in pH units | Parameter 1 | example of value (recommended by Ponsel): 7 pH units
| standard 2 | Calibration standard value (slope) | in pH units | Parameter 1 | example of value (recommended by Ponsel): 4 pH units
| standard 3 | Offset |  | Parameter 2 | 0 mV is measured electronically, so sensor may remain exposed to air.
|  | 0x01 |  |  | - '0' activated electronically
|  | 0x02 |  |  | - Offset calculation and '0' disabled electronically
|  | 0x03 |  |  | - Disablement of '0' only
| standard 4 | Calibration standard value (slope) | in mV | Parameter 2 | example of value (recommended by Ponsel): 240 mV
| standard 5 | NU
| standard 6 | NU
| standard 7 | NU
| standard 8 | NU
| standard 9 | NU
| standard 10 | NU
| standard 11 | NU
| standard 12 | NU
| standard 13 | NU
| standard 14 | NU
| standard 15 | NU
| standard 16 | NU

## ORP Annular
| POD Description | "ORP/Redox/Temperature PONSEL " |  |  |  |  | Information available
| Default address | 21 |  |  |  |  | Used for measurement parameter setting
| Serial number | SN-PORPx-xxxx |  |  |  |  | Used for measurement
| Default averaging value | 1 |  |  |  |  | Used for POD parameter setting
Used for parameter calibration
System, used exclusively by Ponsel
Not used
|  | Temperature measurement | Parameter 1 | Parameter 2 | Parameter 3 | Parameter 4
| Parameter measurement : | Temperature | NU | Redox | NU | NU
| Unit : | °C | NU | in mV | NU | NU
| Managed by history or by site : | historical | NU | historical | NU | NU
| Compensates Temperature returned by Master : | NU | NU | NU | NU | NU
| Compensates 1 returned by Master : | NU
| Compensates 2 returned by Master : | NU
| Compensates 3 returned by Master : | NU
Measurement configuration : |  | In use | NU | In use | NU | NU
Measurement type configuration | Differed mode in use : | NO | NU | NO | NU | NU
| Corrected measurement (by default) : | returns measurement | NU | returns measurement | NU | NU
| Corrected measurement 1 : | NU
| Corrected measurement 2 : | NU
| Gross measurement : | NU
| Range configuration : CgGamme | NU
Measurement status : |  | In use | NU | In use | NU | NU
Measurement status | Measurement OK with INFO 1 | NU | NU | NU | NU | NU
| Measurement OK with INFO 2 | NU | NU | NU | NU | NU
| Measurement impossible with INFO 3 | NU | NU | NU | NU | NU
| Measurement impossible with INFO 4 | NU | NU | NU | NU | NU
Description of calibration coefficients : all coefficients are based on default coefficient.
|  | Information : | Unit : | Coefficient used for parameter : | Comments :
| Offset Temp Coefficient | Temperature offset | °C | Temperature
| Temp Slope Coefficient | Temperature slope | % | Temperature
| Coef 1 | NU
| Coef 2 | NU
| Coef 3 | Redox offset | in mV | Parameter 2
| Coef 4 | Redox slope | % | Parameter 2
| Coef 5 | NU
| Coef 6 | NU
| Coef 7 | NU
| Coef 8 | NU
| Coef 9 | NU
| Coef 10 | NU
| Coef 11 | NU
| Coef 12 | NU
| Coef 13 | NU
| Coef 14 | NU
| Coef 15 | NU
| Coef 16 | NU
Description of calibration standards :
|  | Master returns : | Unit : | Calibration standard used for parameter : | Comments :
| Calibration standard 1 for temperature : | Calibration standard value (offset) | °C | Temperature | Value recommended by Ponsel : near 0°C
| Calibration standard 2 for temperature : | Calibration standard value (slope) | °C | Temperature | Value recommended by Ponsel : near 25°C
| Standard 1 | NU
| Standard 2 | NU
| Standard 3 | Offset |  | Parameter 2 | 0 mV is measured electronically, so sensor may remain exposed to air.
|  | 0x01 |  |  | - '0' activated electronically
|  | 0x02 |  |  | - Offset calculation and '0' disabled electronically
|  | 0x03 |  |  | - Disablement of '0' only
| Standard 4 | Calibration standard value (slope) | in mV | Parameter 2 | Example of value (recommended by Ponsel) : 240 mV
| Standard 5 | NU
| Standard 6 | NU
| Standard 7 | NU
| Standard 8 | NU
| Standard 9 | NU
| Standard 10 | NU
| Standard 11 | NU
| Standard 12 | NU
| Standard 13 | NU
| Standard 14 | NU
| Standard 15 | NU
| Standard 16 | NU
u16 variables in RAM (0x006F to 0x007E)
| Float variables in RAM (0x007F to 0x009D) | [5 empty] | Variable 1 | NU
| Variable 1 | NU |  |  |  |  | Variable 2 | NU
| Variable 2 | NU |  |  |  |  | Variable 3 | NU
| Variable 3 | NU |  |  |  |  | Variable 4 | NU
| Variable 4 | NU |  |  |  |  | Variable 5 | NU
| Variable 5 | NU |  |  |  |  | Variable 6 | NU
| Variable 6 | NU |  |  |  |  | Variable 7 | NU
| Variable 7 | NU |  |  |  |  | Variable 8 | NU
| Variable 8 | NU |  |  |  |  | Variable 9 | NU
| Variable 9 | NU |  |  |  |  | Variable 10 | NU
| Variable 10 | NU |  |  |  |  | Variable 11 | NU
| Variable 11 | NU |  |  |  |  | Variable 12 | NU
| Variable 12 | NU |  |  |  |  | Variable 13 | NU
| Variable 13 | NU |  |  |  |  | Variable 14 | NU
| Variable 14 | NU |  |  |  |  | Variable 15 | NU
| Variable 15 | NU |  |  |  |  | Variable 16 | NU
| Variable 16 | NU

## Conductivity-C4E
| POD description | "C4E /Salinite/Temp PONSEL " | [5 empty] | Information available
| default address | 30 | [5 empty] | Used for measurement parameter setting
| serial number | SN-PC4EA-xxxx | [5 empty] | Used for measurement:
| Default averaging value | 10 | [5 empty] | Used for POD parameter setting:
Used for parameter calibration
System, used exclusively by Ponsel
[7 empty] | NU | Not used
|  | Temperature measurement | Parameter 1 | Parameter 2 | Parameter 3 | Parameter 4
| Parameter measurement: | Temperature | C4E conductivity | Salinity | TDS - Kcl | NU
| Unit: | °C | µS/cm | ppt | ppm | NU
| Managed by history or by site: | historical | historical | historical | historical | NU
| Compensates Temperature returned by Master | NU | used in °C |  |  | NU
| Compensates 1 returned by Master | NU
| Compensates 2 returned by Master | NU
| Compensates 3 returned by Master | NU
Measurement configuration: |  | Used | Used for parameters 1, 2 and 3 | NU | NU | NU
Measurement type configuration | differed Mode used : | NO | NO | NU | NU | NU
| corrected measurement (default value): | returns measurement | returns measurement | NU | NU | NU
| corrected measurement 1: | NU
| corrected measurement 2: | NU
| gross measurement: | NU
| Range configuration: CgGamme | NU | Use of 4 ranges for Parameters 1, 2 and 3: - range 1: 0 / 200 µS/cm - range 2: 0 / 2000 µS/cm - range 3: 0 / 20.000 µS/cm - range 4: 0 / 200.000 µS/cm | NU | NU | NU
Measurement status: |  | In use | In use | NU | NU | NU
Measurement status | Measurement OK with INFO 1 | NU | "Slight" oscillation on measurement | NU | NU | NU
| Measurement OK with INFO 2 | NU | NU | NU | NU | NU
| Measurement impossible with INFO 3 | NU | NU | NU | NU | NU
| Measurement impossible with INFO 4 | NU | NU | NU | NU | NU
Description of calibration coefficients : all coefficients are based on default coefficient.
| Note: | Salt content and TDS are not calibrated: these parameters are derived from the conductivity measurement.
|  | Information: | Unit: | Coefficient used for parameter: | Comments:
| Offset Temp Coefficient | Temperature offset | °C | Temperature
| Temp Slope Coefficient | Temperature slope | % | Temperature
| Coef 1 | Range offset 1 | µS/cm | Parameters 1, 2 and 3.
| Coef 2 | Range slope 1 | % | Parameters 1, 2 and 3.
| Coef 3 | Range offset 2 | µS/cm | Parameters 1, 2 and 3.
| Coef 4 | Range slope 2 | % | Parameters 1, 2 and 3.
| Coef 5 | Range offset 3 | µS/cm | Parameters 1, 2 and 3.
| Coef 6 | Range slope 3 | % | Parameters 1, 2 and 3.
| Coef 7 | Range offset 4 | µS/cm | Parameters 1, 2 and 3.
| Coef 8 | Range slope 4 | % | Parameters 1, 2 and 3.
| Coef 9 | NU
| Coef 10 | NU
| Coef 11 | NU
| Coef 12 | NU
| Coef 13 | NU
| Coef 14 | NU
| Coef 15 | NU
| Coef 16 | NU
Description of calibration standards :
|  | Master returns: | Unit: | Calibration standard used for parameter: | Comments:
| calibration standard 1 for temperature | Calibration standard value (offset): | °C | Temperature | value recommended by Ponsel: near 0°C
| calibration standard 2 for temperature | Calibration standard value (slope) | °C | Temperature | value recommended by Ponsel: near 25°C
| standard 1 | 0' value (offset) | in µS/cm | Parameters 1, 2 and 3. | Ponsel recommends that the sensor be placed in the air.
| standard 2 | Calibration standard value (slope) | in µS/cm | Parameters 1, 2 and 3. | example of value (recommended by Ponsel): 84 µS/cm
| standard 3 | 0' value (offset) | in µS/cm | Parameters 1, 2 and 3. | Ponsel recommends that the sensor be exposed to air.
| standard 4 | Calibration standard value (slope) | in µS/cm | Parameters 1, 2 and 3. | example of value (recommended by Ponsel): 1413 µS/cm
| standard 5 | 0' value (offset) | in µS/cm | Parameters 1, 2 and 3. | Ponsel recommends that the sensor be exposed to air.
| standard 6 | Calibration standard value (slope) | in µS/cm | Parameters 1, 2 and 3. | example of value (recommended by Ponsel): 12880 µS/cm
| standard 7 | 0' value (offset) | in µS/cm | Parameters 1, 2 and 3. | Ponsel recommends that the sensor be exposed to air.
| standard 8 | Calibration standard value (slope) | in µS/cm | Parameters 1, 2 and 3. | example of value (recommended by Ponsel): 111800 µS/cm
| standard 9 | NU
| standard 10 | NU
| standard 11 | NU
| standard 12 | NU
| standard 13 | NU
| standard 14 | NU
| standard 15 | NU
| standard 16 | NU

## Conductivity CTZ
| POD description | "CTZ/Salinity/Temp PONSEL " | [5 empty] | Information available
| default address | 50 | [5 empty] | Used for measurement parameter setting
| serial number | SN-PCTZA-xxxx | [5 empty] | Used for measurement :
| Default averaging value | 10 | [5 empty] | Used for POD parameter setting :
Used for parameter calibration
System, used exclusively by Ponsel
[7 empty] | NU | Not used
|  | Temperature measurement | Parameter 1 | Parameter 2 | Parameter 3 | Parameter 4
| Parameter measurement: | Temperature | Conductivity CTZ | Salinity | Conductivity CTZ without temperature compensation | NU
| Unit: | °C | mS / cm | g/Kg | mS / cm | NU
| Managed by history or by site: | history | history | history | history | NU
| Compensates Temperature returned by Master | NU | used in °C |  | NU | NU
| Compensates 1 returned by Master | NU | Alpha en %/°C Value by default (power of the Pod : 0 %/°C --> Uncompensated measurement).
| Compensates 2 returned by Master | NU
| Compensates 3 returned by Master | NU
Config of mesure : |  | Used | used for the parameters 1, 2 and 3 | NU | NU | NU
Configuring the nature of the measurements | Batch mode used : | NO | NO | NU | NU | NU
| Measurement rectified (by default) : | send the measurement back | send the measurement back | NU | NU | NU
| Rectified measurement 2 : | NU
| Rectified measurement 2 : | NU
| Gross measurement : | NU
| Range Configuration : Cg range | NU | The use of 2 ranges ofr parameter 1, 2 and 3: (no automatic range) - range 1: 0 / 100 mS/cm - range 2 : 0 / 700 and + mS/cm | NU | NU | NU
Measurement status : |  | Used | Used | Used | Used | NU
Measurement status | Measurement OK with the INFO 1 | The supply voltage (Vcc_micro) is too low. |  |  |  | NU
| Measurement OK with the INFO 2 | NU | NU | NU | NU | NU
| Measurement is impossible with the INFO 3 | The supply voltage (Vcc_micro) is too bad. |  |  |  | NU
| Measurement is impossible with the INFO 4 | NU | NU | NU | NU | NU
Descriptions of calibration coefficients : all coefficients are relative to factory coef
| Comments : | Salinity does not calibrate these parameters are derived from the measurement of conductivity.
|  | Information | Unité : | Coeff servant au paramétre : | Remarques :
| Coef Offset Temp | Offset of temperature | °C | Temperature
| Coef Pente Temp | Slope of temperature | % | Temperature
| Coef 1 | Offset of range 1 | mS/cm | Parameter 1 and 2.
| Coef 2 | Slope of range 1 | % | Parameter 1 and 2.
| Coef 3 | Offset of range 2 | mS/cm | Parameter 1 and 2.
| Coef 4 | Slope of range 2 | % | Parameter 1 and 2.
| Coef 5 | NU
| Coef 6 | NU
| Coef 7 | NU
| Coef 8 | NU
| Coef 9 | NU
| Coef 10 | NU
| Coef 11 | NU
| Coef 12 | NU
| Coef 13 | NU
| Coef 14 | NU
| Coef 15 | NU
| Coef 16 | NU
Description of measure :
|  | The master returns : | Unit | Cal for parameter | Comments :
| Cal 1 of temperature | The value of measure (offset) | °C | Temperature | Value recommended by Ponsel : close to 0°C
| Cal 2 of temperature | The value of measure (slope) | °C | Temperature | Value recommended by Ponsel : close to 25°C
| Cal 1 | The value of measure, close to '0 mS/cm' (offset) | unit in mS/cm | Parameter 1 and 2. | Ponsel recommends to put the sensor in demineralized water
| Cal 2 | slope |  | Parameter 1 and 2. | 2 kinds of calibration for the slope : - Manual status : Put the sensor in buffer and send the buffer value in mS/cm (value recommended by Ponsel : XXX mS/cm). - Automatic calibration : Put the sensor in the air. |  |  |  | The automatic calibration is not currently available on our pods
|  | the value of measure | unit in mS/cm |  | - Calculating the slope (Manual calibrationl)
|  | 0x01 |  |  | - Activation of the automatic calibration. (simulation of conductivity in….mS/cm)
|  | 0x02 |  |  | - Calculating the slope and et deactivation of the automatic calibration
|  | 0x03 |  |  | - Deactivation of the automatic calibration only
| Cal 3 | the value of measure, close to '0mS/cm' (offset) | unit in mS/cm | Parameter 1 and 2. | Ponsel recommends to put the sensor in demineralized water
| Cal 4 | slope |  | Parameter 1 and 2. | 2 kinds of calibration for the slope : - Manual status : Put the sensor in buffer and send the buffer value in mS/cm (value recommended by Ponsel : XXX mS/cm). - Automatic calibration : Put the sensor in the air. |  |  |  | The automatic calibration is not currently available on our pods
|  | the value of measure | unit in mS/cm |  | - Calculating the slope (Manual calibrationl)
|  | 0x01 |  |  | - Activation of the automatic calibration. (simulation of conductivity in….mS/cm)
|  | 0x02 |  |  | - Calculating the slope and et deactivation of the automatic calibration
|  | 0x03 |  |  | - Deactivation of the automatic calibration only
| Cal 5 | NU
| Cal 6 | NU
| Cal 7 | NU
| Cal 8 | NU
| Cal 9 | NU
| Cal 10 | NU
| Cal 11 | NU
| Cal 12 | NU
| Cal 13 | NU
| Cal 14 | NU
| Cal 15 | NU
| Cal 16 | NU

## Turbidity Suspended solid
| POD Description | "TU/MES/VB 5 PONSEL " "VB 5 PONSEL " (parameter 2 and 3 not available) | [5 empty] | Information available
| Default address | 60 | [5 empty] | Used for measurement parameter setting
| Serial number | SN-PTURx-xxxx | [5 empty] | Used for measurement
| Default averaging value | 10 | [5 empty] | Used for POD parameter setting
Used for parameter calibration
System, used exclusively by Ponsel
[7 empty] | NU | Not used
|  | Temperature measurement | Parameter 1 | Parameter 2 | Parameter 3 | Parameter 4
| Parameter measurement : | Temperature | Sludge blanket | MES (Suspended Solid) | Turbidity | NU
| Unit : | °C | % | g/l | FAU | NU
| Managed by history or by site : | historical | historical | Site | historical | NU
| Compensates Temperature returned by Master : | NU | NU | NU | NU | NU
| Compensates 1 returned by Master : | NU
| Compensates 2 returned by Master : | NU
| Compensates 3 returned by Master : | NU
Measurement configuration : |  | In use | In use | In use | In use | NU
Measurement type configuration | Differed mode in use : | NO | NO | NO | NO | NU
| Corrected measurement (by default) : | returns measurement | returns measurement | returns measurement | returns measurement | NU
| Corrected measurement 1 : | NU
| Corrected measurement 2 : | NU
| Gross measurement : | NU
| Range configuration : CgGamme | NU | NU | NU | NU | NU
Measurement status : |  | In use | In use | In use | In use | NU
Measurement status | Measurement OK with INFO 1 | Power supply is too low or too loud |  |  |  | NU
| Measurement OK with INFO 2 | NU | There is too much exterior light |  |  | NU
| Measurement impossible with INFO 3 | Power supply is too low or too loud or there is at least 1 defective peripheral |  |  |  | NU
| Measurement impossible with INFO 4 | NU | There is too much exterior light |  |  | NU
Description of calibration coefficients : all coefficients are based on default coefficient.
|  | Information : | Unit : | Coefficient used for parameter : | Comments :
| Offset Temp Coefficient | Temperature offset | °C | Temperature
| Temp Slope Coefficient | Temperature slope | % | Temperature
| Coef 1 | FAU Offset | in ua | Parameter 3
| Coef 2 | FAU Slope | % | Parameter 3
| Coef 3 | VB Offset | in ua | Parameter 1
| Coef 4 | MES Offset (Suspended Solid) | in ua | Parameter 2
| Coef 5 | MES Slope (Suspended Solid) | % | Parameter 2
| Coef 6 |  |  |  | Used to memorise substance value
| Coef 7 |  |  |  | Used to memorise dry weight value
| Coef 8 | NU
| Coef 9 | NU
| Coef 10 | NU
| Coef 11 | NU
| Coef 12 | NU
| Coef 13 | NU
| Coef 14 | NU
| Coef 15 | NU
| Coef 16 | NU
Description of calibration standards :
|  | Master returns : | Unit : | Calibration standard used for parameter : | Comments :
| Calibration standard 1 for temperature : | Calibration standard value (offset) | °C | Temperature | Value recommended by Ponsel : near 0°C
| Calibration standard 2 for temperature : | Calibration standard value (slope) | °C | Temperature | Value recommended by Ponsel : near 25°C
| Standard 1 | Calibration standard value, near '0 FAU' (offset) | in FAU | Parameter 3 | Value recommended by Ponsel : demineralised water 0 FAU
| Standard 2 | Calibration standard value (slope) | in FAU | Parameter 3 | Value recommended by Ponsel : 2000 FAU (mid-range)
| Standard 3 | Calibration standard value | unité du VB en % | Parameter 1 | Value recommended by Ponsel : demineralised water, let 100%
| Standard 4 | Calibration standard value, near '0 g/l' (offset) | in g/l | Parameter 2 | Value recommended by Ponsel : demineralised water 0 g/l
| Standard 5 | Calibration standard value (slope) | in g/l | Parameter 2
| Standard 6 |  |  |  | Used to memorise substance value
| Standard 7 |  |  |  | Used to memorise dry weight value
| Standard 8 | NU
| Standard 9 | NU
| Standard 10 | NU
| Standard 11 | NU
| Standard 12 | NU
| Standard 13 | NU
| Standard 14 | NU
| Standard 15 | NU
| Standard 16 | NU